# Title Page

Supplement for “Technical Note: A new method to evaluate fluoroscopic system collimator performance”

Authors:

Joseph R. Steiner1*, Courtney K. Morrison2, Mayur Vaya2, Nicholas Bevins3, Jeremy Christophel2, Matt Vanderhoek2

1Department of Radiology, University of Chicago, Chicago, IL 60637

2Department of Diagnostic Radiology and Nuclear Medicine, RUSH University Medical Center, Chicago IL 60612

2Department of Radiology, Henry Ford Health, Detroit, MI 48202

3Department of Radiology, Maine Medical Center, Portland, ME 04102

*Corresponding author: Contact at jrsteiner@uchicago.edu

Running title:

Fluoroscopy Collimator Performance Evaluation

# Supplement Introduction

This document provides supplemental information regarding the manuscript: “Technical Note: A new method to evaluate fluoroscopic system collimator performance.” This supplement:

- Clarifies the assumptions made in the manuscript regarding the SFD.

- Evaluates the effect of small inaccuracies in the measurement of the SFD.

- Analyzes the uncertainty inherent to the measurement process.

# Supplement Section 1: Clarification of assumptions

In the manuscript, it was assumed the SFD was the minimum SSD for three reasons. The first reason is that the SSD will always be less than the actual distance of the source to the rulers/film. Thus, EW and EL (and similarly EW+EL) computed using Equations 2 and 3 with SFD=SSD will be larger than if the true SFD was used; this is conservatively safe as EW and EL will always be overestimated. The second reason is that the orthogonal rulers used in the work are overlapping; the SFD is different for one ruler relative to the other. Choosing a conservatively safe approach that requires only one distance measurement (minimum SSD) greatly simplifies the procedure and reduces the likelihood of data input error. As described in Section 2, using a measured SFD greater than the actual SFD will underestimate EW, EL and EW + EL which can result in a system with EW and/or EL > 3%SID and/or EW +EL > 4%SID appearing to satisfy federal regulation. The third reason is that federal regulation specifies minimum SSD for fluoroscopy systems; measuring the minimum SSD satisfies this regulation.

A test case using typical geometry for a mini C-arm using the method described in the manuscript is used to demonstrate the effect of this assumption. For the orthogonal rulers used in this work, each arm has a thickness of 3 mm and the arms are separated by a 2 mm gap. In the worst case scenario, for one ruler arm SFD=SSD + 3 mm and for the other ruler arm the SFD=SSD + 8 mm. In Figure S.1, the excess distances of 3 mm and 8 mm are represented as Δz. To compute EW and EL, IF and IDI are measured along ±x and ±y, and the SSD was measured.

Figure S.1: Geometry showing the source plane, the SSD, the SFD, and Δz.

Equation S.1 can be used to calculate the overestimation EWOE of the EW, assuming SFD equals SSD (the assumption) and SSD + Δz where Δz=3 mm or 8 mm and EW is defined in Equation 2. The formalism for computing the overestimation ELOE of the EL is the same, noting that if EWOE and ELOE are computed from a single test case, each will use a different Δz. For the purposes of this example, the minimum limits of SSD specified in Federal Regulation 21CFR1020.32(g)(1,2) were utilized and are reported in Table S.1. The actual measured minimum SSD is generally larger than these federal limits.

EWOE=EWSSDEWSSD+Δz-1×100%=IF,x-IF, -x-IDI,x-IDI, -xSSDIF,x-IF, -x-IDI,x-IDI, -xSSD+Δz-1×100%=ΔzSSD×100% #Eq. S.1

| Table S.1: Percent by which EW and EL are increased by assuming the rulers/film occupy plane of the SSD instead of SSD+Δz. |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| System Type | Minimum SSD (cm) | EW |  | EL |  |  |
|  |  | Δz (cm) | EWOE | Δz (cm) | ELOE |  |
| Fixed C-arm | 20 | 0.3 | 1.5% | 0.8 | 4.0% |  |
| Mini C-arm (non-surgical) | 19 | 0.3 | 1.6% | 0.8 | 4.2% |  |
| Mini C-arm (surgical) | 10 | 0.3 | 3.0% | 0.8 | 8.0% |  |
| Mobile C-arm (non-surgical) | 30 | 0.3 | 1.0% | 0.8 | 2.7% |  |
| Mobile C-arm (surgical) | 20 | 0.3 | 1.5% | 0.8 | 4.0% |  |
| Rad/fluoro (non-surgical) | 38 | 0.3 | 0.8% | 0.8 | 2.1% |  |
| Rad/fluoro (surgical) | 20 | 0.3 | 1.5% | 0.8 | 4.0% |  |

The results of this calculation demonstrate that the values of EW and EL will be overestimated by 0.8% to 8.0% depending on the ruler orientation and the minimum SSD, which depends on system type.  An overestimate of EW or EL by 8.0% is substantial and may result in false positives of collimators failing their performance evaluation; if EW or EL is calculated as 3%SID (limit for failure) using this assumption, the true EW or EL would be 2.76%SID which would satisfy federal regulation. For cases of failure when using this assumption, EW and EL can be more accurately calculated using careful measurement of the SFD for each ruler arm to evaluate if the system needs service. This reduces the complexity for most systems that comply while providing the flexibility to perform more accurate measurements and calculations if needed.

# Supplement Section 2: The effect of small inaccuracies in the measurement of the SFD

As shown in Supplement Section 1, measuring the SFD as 3 mm or 8 mm less than the true SFD resulted in overestimations of EW and EL ranging from 0.8% to 8.0%. By measuring the SFD to the actual film locations, a more accurate value of EW and EL (and the sum of those terms) can be achieved. However, caution must be exercised when measuring to the rulers, because if the SFD measured (SFDmeas) and used for calculation is larger than the true SFD (SFDtrue), EW and EL will be underestimated, which may result in not identifying a fluoroscopy system collimator that does not satisfy the limits specified in Federal Regulation 21CFR1020.32. As the amount of under- or overestimation of EW and EL will be the same for a given SFDtrue and SFDmeas, two new variables are defined; ΔSFD=SFDtrue-SFDmeas (which is equivalent to Δz in Section 1) and ED is the excess dimension which represents EW and EL;  ED is calculated using Equation S.2 (only the calculation along x is shown). Note that Equation S.2 is the same as Equation S.1 where SFDmeas=SSD and SFDtrue=SSD+Δz, with a key difference being that in Equation S.1,  SFDmeas will always be less than or equal to SFDtrue because the rulers and film cannot be located nearer to the source than the minimum SSD.

ED=EDSFDmeasEDSFDtrue-1×100%=IF,x-IF, -x-IDI,x-IDI, -xSFDmeasIF,x-IF, -x-IDI,x-IDI, -xSFDtrue-1×100%=SFDtrueSFDmeas-1×100% #Eq. S.2

Figure S.2 shows the under- and overestimation for ΔSFD ranging from -2 cm to 2 cm for SFDtrue=11.5 cm, 21.0 cm, and 31.0 cm, which are representative of mini-C-arms, surgical fluoroscopy systems, and general fluoroscopy systems. The same trends will occur with any SFDtrue.

Figure S.2: Measured ED over-and underestimation compared to the true ED as a function of ΔSFD for SFDtrue=11.5 cm,21.0 cm,and 31.0 cm.

As seen in Figure S.2, if ΔSFD is negative, the calculated ED will be less than the actual ED. In the worst case scenario shown by the chart, if ΔSFD = - 2.0 cm, this underestimation will be -17.4%, -9.5%, and -6.5% for SFDtrue=11.5 cm, 21.0 cm, and 31.0 cm, respectively. This can result in false negatives, where it appears the fluoroscopy system collimator performance satisfies federal regulation but in fact does not. This is the most important limitation of this method and extreme caution must be exercised if SFDmeas is utilized for calculation of EW and EL.

# Supplement Section 3: Analysis of measurement uncertainty

Finally, in the manuscript that two uncertainties were reported. The first uncertainty is σ (%SID), which the typical standard deviation of the mean value of EW, EL, and EW + EL for each sample of measurements. σ (%SID) was reported solely as a comparative metric as it is unlikely the values of EW, EL, and EW + EL are normal, as they are dependent on the state or condition of each fluoroscopy system collimator assessed. While other comparative metrics could be used (e.g., interquartile ranges), the value again is likely limited and the inclusion of this information may detract from the purpose of this manuscript, which is to outline a robust and simple method for evaluating fluoroscopy collimator performance.

The second uncertainty is the propagation of error from the measurement process. Reproducing Equation 2 and 3 from the manuscript as Equation S.3 and S.4 respectively

EW=IF,x-IF, -x-IDI,x-IDI, -xSFD×100%#Eq. S.3

EL=IF,y-IF, -y-IDI,y-IDI, -ySFD×100%#Eq. S.4

it is clear that if SFD has a constant measurement uncertainty (e.g., 0.5 mm), the uncertainty associated with the measurement will be larger if the rulers/film are located on the x-ray exit aperture (≈SSD) relative to the image receptor cover (≈SID); this is described below. Importantly, the analysis presented below is simplified to encompass only expected measurement error. The systemic errors associated with the assumption SFD = SSD have been thoroughly described in Section 1 and Section 2 and are not included in this section.

To determine the amount of error associated with measurement, it is assumed all measured distances have an inherent uncertainty of 0.5 mm. This is because the radiopaque rulers used have graduations of 1 mm and the measurement of the minimum SSD was measured and/or reported with an accuracy of 1 mm. Using typical propagation of error, the measurement uncertainty σEW in EW is calculated using the uncertainty of each measurement (indicated by σ with the measurement variable as a subscript) as follows, noting the calculation of EL will be equivalent. First, Equation S.3 can be rewritten as

EW=IF,x-IF, -x-IDI,x+IDI, -xSFD×100%

which, if all uncertainties of all terms in the numerator are assumed to be the same (σI), results in an uncertainty of the numerator being

σnum=σIF,x2+σIF,-x2+σIDI,x2+σIDI,-x2=4σI2=2σI

The uncertainty of the denominator SFD is simply σSFD.

Estimation of the uncertainty of EW (σEW) is problematic because is it very likely that the numerator IF,x-IF, -x-IDI,x+IDI, -x will be close to or equal to zero; thus the following methodology is used: A given function z=zx,y has the associated uncertainty σz=δzδx⋅σx2+δzδx⋅σy2.  If z=xy, the partial derivatives are δzδx=1y and δzδy=-xy2. With substitution

σz=σxy2+-xσyy22

Using this derivation with the numerator and denominator for EW as defined above,

σEW=2σISFD2+-IF,x-IF, -x-IDI,x+IDI, -x×σSFDSFD22×100%

and thus the uncertainty associated with measurement is dependent upon the SFD, where a smaller SFD results in a larger relative uncertainty as %SID.

As a representative example, consider a mini-C-arm with SFD=11.5 cm,  IF,x-IF, -x-IDI,y-IDI, -y=0.25 cm, and SID = 44 cm. EW and σEW are

EW=IF,x-IF, -x-IDI,x+IDI, -xSFD×100%=2.5 mm115 mm×100%=2.2%SID

σEW=2*0.5 mm115 mm2+-2.5 mm×0.5 mm115 mm22×100%=0.87%

As a comparative metric, consider if the SFD is equal to SID=440 mm, which is similar to the case where the rulers/film are located on the image receptor cover. The numerator can be calculated as (noting that this is an estimate as it was not measured)

IF,x-IF, -x-IDI,x+IDI, -x=EW×SFD100%=2.2%×440 mm100%=9.68 mm

and from this the uncertainty would be

σEW=2*0.5 mm440 mm2+-9.68 mm×0.5 mm440 mm22×100%=0.27%

which is less than the uncertainty measured at the smaller SFD. This is a limitation of the methodology described in this manuscript; one must decide if the benefits of performing fluoroscopy collimation testing using the methodology in this manuscript outweigh the larger associated measurement uncertainty. Importantly, this measurement uncertainty is generally less than the error if the SFD is not accurately measured (Section 2).

As one final note, the uncertainty of EL is calculated in the same manner as EW, and the uncertainty of EL+EW is calculated as

σEL+EW=σEL2+σEW2

# Supplement Summary

This supplement illustrates the challenges associated with measurement error and uncertainty when assessing collimator performance using the method described in the manuscript. Anyone implementing this method for fluoroscopy collimator performance should note that performing measurements at a smaller SFD will increase measurement uncertainty. However, the error associated with measuring the SFD can result in large variation in the calculated values of EW, EL and EW+EL. Measurement of the SFD is the most critical aspect to ensure accuracy of this method, and the reader may consider assuming SFD=SSD. This assumption can result in more systems failing collimator evaluation; however in these specific cases more careful measurements of the SFD can be performed to obtain more accurate calculated values. If the SFD is overestimated, it is possible that systems failing collimator performance may be missed, and so extreme caution should be exercised when measuring the SFD.