## explanation
This model takes Dual Inlet results, δ13C, δ18O and Δ47, and parameters as input data, calculates expected instrument readings, applies baseline, Δ17 and Δ47 effects, then calculates resulting output results using same or different parameter set
Works on entire data set or single point. Can be modified to work on different groups of data
Can compare/fit to real results--sum squares of differences and use built-in Solver to minimize
"Parameters", "baseline-D17-D47" tabs for information on setting the parameter set values, e.g. factors for calculating 17O from 18O, reference gas, model signals, e.g. 16000 mV on m/z 44, baseline, etc.
"datapoints" tab for both data input (yellow) and output (light purple)
"Full Model" tab for calculations
see calculated data, e.g. mV readings, there
See "Uses" below: this version set up to fit real data,with bottom section incorporating average amplitude differences between sample and working gas, see "datapoints-plots" C30
| Equations used for calculations from Santrock, Studley and Hayes--see box to right: | [6 empty] | Assume stochastic distribution, follow protocol outlined in Santrock, et al.
Model input: | yellow highlighted cells, all in "datapoints-plots" tab, other tabs for reference | [6 empty] | 45R=13R+2*17R
"datapoints" tab: | δ13CVPDB and δ18OVSMOW typed in by user, columns D and E, rest of columns to right are output data (see below) | [6 empty] | 46R=2*18R+2*13R*17R+17R*17R
|  | for single point, "datapoints" tab, line 8 | [5 empty] | with: 17R = k*(18R)λ
|  | for multiple samples, matrix starts on line 10, "datapoints" tab | [5 empty] | rearranging: 13R= 45R-2*17R
| Columns A and B for parameter, baseline, linearity, etc. input: | [6 empty] | Substitute in: 46R=2*18R+ 2*(45R-2*17R)*17R+ 17R*17R
|  | explanations and examples are in "Parameters" and "baseline-D17-D47" tabs | [5 empty] | rearranging: 0= -3*17R2+2*45R*17R+2*18R-46R=x
|  | cell positions for input in "datapoints-plots" tab listed in other tabs | [5 empty] | initial 18R = starting 18R; can use 18R(init) = 46R/2.005
To solve, minimize x by varying 18R: using 18R(new)=18R(old) – x/2.0022
"Parameters" tab: | parameter set(s) for 17O determination, input in Column B and C, suggestions to the right | [6 empty] | Assume stochastic distribution, as above, with:
|  | one for data generation from input, tab "datapoints-plots", cells: A6-10 | [6 empty] | 47R=2*13R*18R+2*17R*18R+13R*17R*17R
|  | one for data output calculation, tab "datapoints-plots", cells: B6-10 | [5 empty] | Amplitude: | m/z 44= 16,000 mV (if 100% m/z 44)
|  | N.B. presumes 17O/16O in VSMOW a function of parameter set used | [5 empty] | Relative amplification: | m/z 44 = 1
| 13C/12C for VPDB, line 11 "Parameters" tab, input in "datapoints-plots" tab, columns A or B, line 8 | [7 empty] | m/z 45 = 100
| 18O/16O for VSMOW, line 10 "Parameters" tab, input in "datapoints-plots" tab, columns A or B, line 7 | [7 empty] | m/z 46 = 333
| ref gas value: δ13CVPDB, δ18OVSMOW, B20 and B21; location in "datapoints-plots" tab, A16 and A17 | [7 empty] | m/z 47 = 3333
ref gas offset values should real value be different from listed value, B22 and B23 (offset not tested); input in "datapoints-plots" tab, B16 and B17
"baseline-D17-D47" tab: | Target value and relative amplification for the different m/z being determined, info around cells B5-8; input, "datapoints-plots", A44 to A49
Offset in signal strength between ref gas and samples, cell B9; location in "datapoints-plots" tab: A42
baseline reading, details near cells B13-16, and/or C for single point, suggestions to the right; input "datapoints-plots" A20-23; single point, B20-23
add to data calculated from input prior to output calculation
Δ47: sample and/or reference--given for calculation, default 0; info around cells B18-19; input "datapoints-plots" A25 (reference) and 26 (sample); single point at B26
Δ17: sample and/or reference--given for calculation, default 0, info around cells B20-21; input "datapoints-plots" tab, A28 (sample) and A29 (reference); single point B28
Δ17 and Δ47 implementation:
δ17OVSMOW calculated from parameter set and Δ17 input effect on initial 17R calculation from δ18O input
Δ47 input modifies m47 signal, effects δ47 and Δ47 , effect of excess on m/z 45 and 46 also determined--excess Δ47 means less 13C and 18O for m/z 45 and 46
treating Δ47 as solely due to 13C-18O, and not accounting for trivial effects on 16O
Linearity--on all cups
per shift in 45R, 46R, 47R and in m/z 44 response with increase in signal; type it in as per mil per volt.
typically measured by increasing gas pressure and looking at per shifts vs voltage on m/z 44; here calculated on change in voltage from ref gas values per m/z
The change in response on m/z 44 shows up as apparent linearity in the other cups if amplitude changed from ref gas.
Isotopologues: relative absolute abundances m/z 45, 46 and 47
All isotopologues at a given m/z value are generally assumed to have the same absolute abundances in calculations.
Type in the relative fractional differences between the absolute abundances for the main isotopologue and the major minor one, info around cells B32-34; input "datapoints-plots" tab, A37-39
A value of 1 means all are the same; 0.99 means the minor one has an absolute abundance that is 99% of the major one; 1.01 means the minor one is at 101% of the major one.
|  |  | m/z | main isotopologue | minor, largest one | amount % of signal from minor component, approximate, natural abundance levels | other components--not adjusted
|  |  | m/z 45 | 16O13C16O | 17O12C16O | 6.6
|  |  | m/z 46 | 18O12C16O | 17O13C16O | 0.2 | 17O12C17O
|  |  | m/z 47 | 18O13C16O | 17O12C18O | 3.4 | 17O13C17O
Measured Δ47 if available can be compared to calculated output. Need to modify spreadsheet to use "Solver" function to find best fit by modifying parameters. Parameters that are changing and the best fit results have to be in the same tab.
Does not model fractionation factors, but can be modified as needed
Model output:
"datapoints" tab: | Calculated δ13CVPDB and δ18OVSMOW: using parameter sets from "Parameters" tab, after other modifications from "baseline-D17-D47" applied
δ17OVSMOW calculated from calculated δ18OVSMOW using "Δ47 determination" parameter set in the "Parameters" tab.
δ47 calculated relative to working gas. Currently no definition of a δ47 scale, but approximately: .δ47 = (δ13CVPDB + δ18OVSMOW)sample - (δ13CVPDB + δ18OVSMOW)reference gas
Δ47 also calculated relative to working (reference) gas, "Parameters" tab. Can apply Δ47 offset to reference gas for calculations, "baseline-D17-D47" tab
"Linear fit": plot Δ47 vs δ47, assume a linear relationship, and correct to the line, i.e. de-trend the results
Δ47 slope corr.: subtract out slope effect to detrend the results
Uses:
A: | See effect of increasing m/z 47 baseline shifts on calculated Δ47 results.
B: | See effect of Δ17 on Δ47 results
C: | See effect of different parameter sets, absolute value for 13C/12C ratios and relationship between 17O and 18O, on Δ47 results
Generate data with one parameter set and use another to calculate results shows what happens if samples have one relationship and we use a different one in our calculations
D: | Tease out mixing effects under given set of conditions:
1. Set two end points in 13C and 18O space
2. Pick a point or set of points that correspond to weighted averages of the two end points to represent mixed samples
3. Do the corresponding weighted averages for δ17O and δ47(wg)
4. Compare calculated results to the weighted averages. Differences are approximately Δ17 and Δ47 shifts.
E: | Fit real data:
1. Fit baseline
a. Assumes all runs have the same baseline effect and all data points at same Δ47 and Δ17O
b. Model data has sample and standards running at identical m/z 44.
c. Keep sample and reference baseline values the same.
d. Minimize sum of squares of differences between model output and sample data
i. use Solver
ii. Adjust baseline and sample Δ47 to minimize differences between model output and measured results
e. Modify spreadsheet to fit mulitple groups of samples, e.g. ones with different Δ47, simultaneously.
f. to use modeled baseline for data re-processing, normalize to running m/z 44 amplitudes
2. Fit parameter set
a. Assumes all runs have the same baseline effect and all data points at same Δ47, and implicitly, same Δ17O
b. Model data has sample and standards running at identical m/z 44.
c. Minimize sum of squares of differences between model output and sample data
i. use Solver
ii. Adjust baseline and parameter(s) of interest to minimize differences between model output and measured results
The best fit may not fall at a single point.
Instead of fitting to the baseline, can compare measured data to the de-trended model data (linear fit, Δ47 vs δ47) if baseline is small
3. Fit 17O
a. Assumes all runs have the same baseline effect and all data points at same Δ47, and different readings indicate a non-zero Δ17
b. Model data has sample and standards running at identical m/z 44.
c. Minimize sum of squares of differences between model output and sample data
i. use Solver
ii. Adjust baseline and Δ17O sample of interest to minimize differences between model output and measured result(s).
4. Monitor 17O effects
a. set up measurements so all samples should have same 13C present--only "differences" due to effects of Δ17O.
b. Calculate estimate of Δ17 for different runs.
i. If working with 18O spiked samples, initial estimate is 18O spike does not add any 17O, so difference in 18O from unspiked sample a measure of "missing" 17O.
ii. For Δ17O = 0, 18O to 17O relationship must fit parameter set used for calculating 17O
c. Compare unlabelled runs to spiked runs (either 17O spiked for 18O spiked for apparent loss of 17O)
i. model using measured 18O values
ii. "add in" the effect, i.e. start with unspiked 13C reading and see if model output gives measured 13C reading for spiked sample.
F: | Other
1. Use model output to test data processing algorithm
a. Modify spreadsheet to repeat Δ47 calculation, e.g. use one parameter set to calculate 13C and 18O, and expected values, from model data, then another for D47 calculation
Akin to taking intrument output for 13C and 18O, the processing data for Δ17 using a different parameter set
Can give rise to large Δ45, Δ46 and even initial Δ47 values: but full Δ47 values normal
2. Modify model to account for other potential issues, e.g. linearity (instrument response) effects.
3. etc.
Versions: #5
Version 4 was double adding Δ17, fixed in Version 5. See "full model" columns F and N
Linearity effect added. As per mil per volt change for 45R, 46R and 47R; also for m/z 44
Can now run "samples" offset from working reference in terms of mV. That is, the maximum signal is shifted when calculated mV readings for samples
Changes in response of m/z 44 noticeable when samples are offset in mV from working reference, gives rise to apparent linearity in 45R and 46R if calculated from changes seen, e.g. 13C and 18O, when running at different offsets
Calculated from mV difference from reference for m/z 44, 45, 46 and 47
Difference used to calculate per mil offset in reading
Per mill offset then aplied to the entire mV value, not just the difference.
Large 13C or 18O differences from the reference also gives rise to significant difference in m/z 45 and m/z 46
Linearity effect will also show up here even though m/z 44 mV very close
Relative absolute abundances for isotopologues added.
Different m/z readings comprised of different isotopologues.
Only one isotopologue for m/z 44
Value typed in is for relative absolute abundance for largest minor isotopologue compared to main isotopologue at given m/z
The other minor isotopologues are not affected by the input value, i.e. same absolute abundance as main isotopologue
The minor isotopologues only above 1% of signal only for m/z 45 and 47 at natural abundance levels
If relative absolute abundances set to 1, all are equal; if <1, then minor isotopologue contributes less; more if value >1.
This is for what is in the main CO2 signal, not the breakdown products.
About 10 to 20% of the ions are from the CO2 ions breaking down, giving rise to CO, C and O
Studies on N2O looking at NO formed in the source show apparent differnces in the isotopic ratios in the breakdown product dependent on source conditions.
Westley, M. B., B. N. Popp, and T. M. Rust (2007), The calibration of the intramolecular nitrogen isotope distribution in nitrous oxide measured by isotope ratio mass spectrometry, Rapid Commun. Mass Spectrom., 21(3), 391–405, doi:10.1002/rcm.2828.
Detrending: Δ47 slope corr.: subtract out slope effect to detrend the results Was subtracting out entire line before so values end up around zero, not modeled Δ47 input. That is now corrected
Locked calculations, password to unlock a given sheet: unlock
also lock reporting of results and summary of parameters seen on datapoint-plots tab
input and plots are not locked
example parameter sets are locked too
fixed plots of -10‰ δ13C and +10‰ δ18O lines from data matrix, that is the data sort, line 47 and below in "datapoints-plots" tab fixed
version #5B | Move input to "datapoints-plots" tab
version #6 | In most recent version 5, tried correcting C45 and C46, for D47, but instead of using input for set D47, was using input for D17. Full Model tab, columns Q and R. Swapped input positions, fixed for C47, but not C45 or C46
NOTE: this issue did not affect an older version of #5 where fractionation effects were not yet implemented
new additions: | take input linearity "as measured", i.e. change in 45/44 or 46/44 over change in m/z44 amplitude, and modify it using 44/45 and 44/46 ratios from reference gas so values used in calculations are based shifts in amplitude seen for m/z45 and 46 respectively.
include difference in m/z 44 between working gas and sample, a single shift for all samples.
version #7b-fitting | Normalize PBL input to m/z 44 signal--esp. important when varying relative signal intensity.
"fitting" series
remove section below main input under "datapoints-plots" where input re-sorted on d18O--which helped in some plots
have section for input of D47 from measured runs, "datapoints-plots; J10:J26". From initial data processing, i.e. no subsequent D47 vs d47 correction.
have section for input of m/z 44 differences between working gas and sample, "datapoints-plots; A29:A45"
report differences between measured and model D47 prior to D47 vs d47 correction(s); D51:D67 (top) w/out m/z 44 difference included; E51:E67 (bottom) w/ m/z 44 difference
Range of differnces between model and sample below results
Columns G and H are the squares of the differences, summarized at bottom of results. Minimize one of these values when fitting variable to measured results
3/10/2018 fixed single point baseline issue--single point effect on top of general baseline effect, previously was counting general baseline effect twice
| 2018-03-31 | re-arrange positions for parameter inputs
| 2018-12-01 | minor touch ups

## Parameters
Parameters used to generate data for given sets of δ13CVPDB and δ18OVSMOW values.
Same or different parameter set used to calculate results using reference gas values (CO2 gas, Δ47 calculations)
Focus on clumped isotopes, but may see effect on δ13CVPDB results. | [8 empty] | password to unlock
User input in tab "datapoints-plots", explainations and examples here | [8 empty] | unlock
Light blue for calculations
Type in the parameters below in "datapoints-plots" tab for both generating model data and for calculating results from that set of model data.
cell in datapoints-plots tab
parameters: | data generation | Δ47 determination |  | example parameter sets (cut and paste values only from unshaded column to desired "User Defined" column, initial or solve-D47):
title: | type/copy-paste values into cells |  |  | BAC (2010 IUPAC) | SSH | SSH w/ current 13C | GSR (1995 Gonfiantini) | GSR w/ current 13C/12C | A&B | A&B w/ current 17O
VSMOW, 18O/16O | in "datapoints-plots" tab, |  |  | 0.0020052 | 0.0020052 | 0.0020052 | 0.0020052 | 0.0020052 | A&B | 0.0020052 |  | Gonfiantini:
VPDB, 13C/12C | cells: | cells: |  | 0.01118 | 0.011237 | 0.01118 | 0.011237 | 0.01118 | 0.011237 | 0.011237 |  | VSMOW
lambda | A6-A10 | B6-B10 |  | 0.528 | 0.516 | 0.516 | 0.5164 | 0.5164 | 0.528 | 0.528 |  | 18/16 | 0.0020052
K |  |  |  | 0.010225 | 0.009924 | 0.009924 | 0.009394 | 0.009394 | 0.010277 | 0.010225 |  | 17/16 | 0.0003799
calculated: | calculated--do not overwright |  |  | calculated:
VSMOW, 17O/16O | cell: | cells: |  | 0.00038476 | 0.00040235 | 0.00040235 | 0.00037991 | 0.00037991 | #VALUE! | 0.00038476 |  | avg lambda: | 0.5164
VSMOW, 16O/totalO | A12-A14 | B12-B14 |  | 0.99762 | 0.9976 | 0.9976 | 0.99762 | 0.99762 | #VALUE! | 0.99762 |  | K = 17R/(18R)^lambda
VPDB, 12C/totalC |  |  |  | 0.98894 | 0.98889 | 0.98894 | 0.98889 | 0.98894 | 0.98889 | 0.98889 |  | K= | 0.0093937
Type in the values for the reference (working) gas here |  |  |  | 17O per mil from 2010 | 45.7169 | 45.7169 | -12.5871 | -12.5871 | #VALUE! | 0
reference gas: | cell in datapoints-plots tab |  |  | 17R/13R--note, typically done for both being VPDB-CO2, not VSMOW/VPB
δ13CVPDB | cells: | example: | -3.61 | 0.034414 | 0.035805 | 0.035987 | 0.033809 | 0.033981 | #VALUE! | 0.03424
δ18OVSMOW | A16-A17 |  | 24.99 | difference from 2010 IUPAC in 17R/13R space
| ref gas offset--if true value different from label |  |  |  | 40.42 | 45.7169 | -17.5887 | -12.5871 | #VALUE! | -5.0653
| cells: | per mil, please leave at zero |  |  | 0 to 66 | 80 | 0 to -62 | 50 | 1 | 15 | approx. spread in ppm D47 (after D47 vs d47 corr.)
| B16-B17 | per mil, please leave at zero

## baseline-D17-D47
Generate data for given sets of δ13C VPDB and δ18O VSMOW and reference values (CO2 gas, Δ47 calculations)
User input in tab "datapoints-plots", explainations and examples here
Type in the values for the signal in the main cup, m/z44, and the relative amplifications of the other cups (Duel-Inlet run) in "datapoints-plots" tab
model signals: | cell in datapoints-plots tab | example:
m/z 44, max signal, e.g. mV: | type/copy-paste values into cells | 16000 | value if combine all signals together at amplification factor 1
m/z 45, relative amplification to m/z 44 signal | in "datapoints-plots" tab, | 100 | 44 amplification = 1, e.g. 3e8/3e8; for 45 3e10=300e8 so 300/3, etc.
m/z 46, relative amplification to m/z 44 signal | cells: | 333.3333 | Note: signal can be in any unit as long as used consistently
m/z 47, relative amplification to m/z 44 signal | A44-A49 | 3333.3333
m/z 44 offset: sample from working reference: | cell: A42 | 0 | use in conjunction with linearity, see below
Type in background, e.g. Pressure Baseline, values (0 if none present), Δ47 values and Δ17 as appropriate in "datapoints-plots" tab. Note: have option to modify just one sample.
cell in datapoints-plots tab
modification to signal(s): | all: | single point | example (no effects)
background, m/z 45, e.g. mV | type/copy-paste values into cells |  | 0 | for both sample and reference
background, m/z 46, e.g. mV | in "datapoints-plots" tab, |  | 0 | for both sample and reference
background, m/z 47,e.g. mV | cells: | cells: | 0 | for both sample and reference
background, m/z 44, e.g. mV | A20-A23 | B20-B23 | 0 | for both sample and reference
model Δ47, reference | cells: |  | 0 | use 0 for perfect stochastic distribution, note calculations are normalized to reference, so final reference value will be 0--but other values will shift
model Δ47, sample | A25-A26 | B26 | 0 | applied as part per thousand shift on calc signal, before background adjustment
Δ17, reference, 17O anomaly | cells: |  | 0 | 17O anomaly, sample only, not reference
Δ17, sample, 17O anomaly | A28-A29 | B29 | 0 | 17O anomaly, reference only
N.B. "single point" effect added on to "all"
Type in linearity, in per mil 45R and/or 46R per volt signal, in "datapoints-plots" tab. As measured, using a single gas, change in voltage measured on m/z 44. Here, linearity applied to voltage on m/z 45 or m/z 46 respectively.
cell in datapoints-plots tab
modification to signal(s): | all: |  | example (no effects) |  | example: generate data with IUPAC2010, process with 1995: creates Δ47 vs δ47 slope; after slope correction samples within 11 ppm of each other
linearity, m/z 45, per mil per volt | type/copy-paste values into cells |  | 0 | for both sample and reference | 0.06
linearity, m/z 46, per mil per volt | in "datapoints-plots" tab, |  | 0 | for both sample and reference | 0
linearity, m/z 47, per mil per volt | cells: |  | 0 | for both sample and reference
linearity, m/z 44, per mil per volt* | A31-A34 |  | 0 | for both sample and reference | *shows up when working reference and samples are at different signal strengths, basically gives rise to apparent linearity shifts in the same direction for 45R and 46R
Type in relative absolute abundance for largest minor isotopologue compared to main isotopologue at a given m/z ratio in "datapoints-plots" tab. Mainly affects m/z 45 and 47, does not apply to m/z 44
cell in datapoints-plots tab
modification to signal(s): | all: |  | example (no effects)
linearity, m/z 45, per mil per volt | cells: |  | 1 | for both sample and reference
linearity, m/z 46, per mil per volt | A37-A39 |  | 1 | for both sample and reference
linearity, m/z 47, per mil per volt |  |  | 1 | for both sample and reference
linearity, m/z 44, per mil per volt* | n/a* |  | n/a | for both sample and reference | * not applicable, only one isotopologue present

## EG_data
EG1 |  |  |  | BAC (2010) parameter set processing values: | [7 empty] | GSR (1995) parameter set
|  |  |  | d13C | d18O | D47--not D47 vs d47 slope corrected | m/z44 diff, avg | d47 |  |  |  | d13C | d18O | D47--not D47 vs d47 slope corrected | d47
Full- 2010 data_47-interpo_ 150325_124022_clumped_HG-26-b7M |  |  |  | -40.9243 | 42.3331 | -0.072641 |  | -19.6404 |  |  |  | -40.8563 | 42.3326 | -0.10543 | -19.6404
Full- 2010 data_47-interpo_ 150325_163111_clumped_HG-26-74H |  |  |  | -10.3936 | 75.936 | 0.14778 |  | 43.8742 |  |  |  | -10.3155 | 75.9366 | 0.10983 | 43.8742
Full- 2010 data_47-interpo_ 150325_194920_clumped_EG-26-74L |  |  |  | -10.3582 | 16.6071 | -0.045341 |  | -14.8584 |  |  |  | -10.3618 | 16.6069 | -0.043291 | -14.8584
Full- 2010 data_47-interpo_ 150325_230848_clumped_EG-26-87M-b |  |  |  | -40.9665 | 16.2066 | -0.1249 |  | -44.7543 |  |  |  | -40.9344 | 16.2057 | -0.13953 | -44.7543
Full- 2010 data_47-interpo_ 150326_121142_clumped_EG-26-74L-rerun |  |  |  | -10.356 | 16.6859 | -0.024185 |  | -14.7578 |  |  |  | -10.3595 | 16.6857 | -0.022186 | -14.7578
Full- 2010 data_47-interpo_ 150327_091447_clumped_EG-26-74M-rerun |  |  |  | -10.4409 | 42.6773 | 0.037474 |  | 10.8679 |  |  |  | -10.4085 | 42.6774 | 0.021841 | 10.8679
Full- 2010 data_47-interpo_ 150327_123534_clumped_EG-26-87M-c |  |  |  | -40.6019 | 42.3336 | -0.049625 |  | -19.2995 |  |  |  | -40.5343 | 42.3331 | -0.082228 | -19.2995
Full- 2010 data_47-interpo_ 150327_155039_clumped_EG-26-87H |  |  |  | -40.4442 | 75.6445 | 0.05557 |  | 12.9084 |  |  |  | -40.3309 | 75.6444 | 4.1417e-05 | 12.9084
Full- 2010 data_47-interpo_ 150327_184249_clumped_EG-26-87L |  |  |  | -40.4166 | 16.1462 | -0.15665 |  | -44.314 |  |  |  | -40.3853 | 16.1453 | -0.17092 | -44.314
Full- 2010 data_47-interpo_ 150328_114031_clumped_EG-26-87L-rerun |  |  |  | -40.4239 | 16.1571 | -0.11937 |  | -44.2748 |  |  |  | -40.3925 | 16.1563 | -0.13365 | -44.2748
Full- 2010 data_47-interpo_ 150328_152247_clumped_EG-26-74M-rerun2 |  |  |  | -10.4235 | 42.6207 | 0.030691 |  | 10.8224 |  |  |  | -10.3911 | 42.6208 | 0.015105 | 10.8224
2010 parameter set processing values:
EG2 |  |  |  | d13C | d18O | D47 | m/z44 diff, avg | d47
[7 empty] | avg diff: |  |  |  | linearity, 45, 46
151109 | test2 | EG26-1787B-mixed waters |  | -40.6937 | 12.4181 | 0.079755 | 73.9488 | -47.925 | mix runs, D17 anomaly |  | 0.010878 | -40.6673 | 12.4172 | 0.067928 | -47.925
151109 | test3 | EG26-1774B-mixed waters |  | -10.3977 | 12.7577 | 0.035884 | 5.4101 | -18.6134 | data uncorrected |  | 0.056455 | -10.4066 | 12.7573 | 0.040528 | -18.6134
151109 | test4 | EG26-1787A-vsmow2 |  | -40.5562 | 40.6531 | 0.068948 | 41.7444 | -20.7489 | expected correction ca: |  |  | -40.4909 | 40.6526 | 0.037534 | -20.7489
151109 | test5 | EG26-1774A-slap |  | -10.3944 | -16.2288 | 0.038395 | -15.2447 | -47.1873 | subtract 0.02 from 13C and 0.01 from D47 |  |  | -10.4435 | -16.2296 | 0.062808 | -47.1873
151109 | test6 | EG26-1774A-vsmow2 |  | -10.3079 | 40.6721 | -0.01711 | 1.8245 | 8.9646 |  |  |  | -10.2784 | 40.6722 | -0.031316 | 8.9646
151110 | test1 | EG26-1787A-slap 16V |  | -40.7391 | -16.1711 | 0.1275 | -9.5881 | -75.2923 |  |  |  | -40.7522 | -16.1723 | 0.13571 | -75.2923
2010 parameter set processing values:
2010, EG3a | model-B-2010 | model-L-2010 |  | d13C | d18O | D47 | m/z44 diff, avg |  | again D47 values without D47 vs d47 slope correction
170803 |  | 231-cleaned |  | -10.3304 | 43.1263 | 0.0065582 | 259.0365 |  |  |  |  | -10.2975 | 43.1264 | -0.0093165 | 11.3889
170807 |  | 242-26 |  | -10.4303 | 71.5065 | -0.017173 | -306.8653 |  |  |  |  | -10.3583 | 71.507 | -0.052173 | 39.289
170807 |  | 242-26-newref |  | -10.4179 | 71.5307 | -0.022946 | 52.7954 |  |  |  |  | -10.3459 | 71.5312 | -0.057955 | 39.3195
170807 |  | 242-26-rerun |  | -10.4184 | 71.5387 | -0.015852 | 98.6027 |  |  |  |  | -10.3463 | 71.5392 | -0.050866 | 39.3343
170811 |  | 232-26-r1 |  | -10.4221 | 43.4079 | -0.015527 | 8.8688 |  |  |  |  | -10.3887 | 43.408 | -0.031642 | 11.554
170811 |  | 232-26-rr |  | -10.4223 | 43.414 | 0.0044056 | 71.355 |  |  |  |  | -10.3888 | 43.4141 | -0.011714 | 11.5801
170812 |  | 122-26 |  | -40.9799 | 15.2746 | 0.00039531 | 82.4838 |  |  |  |  | -40.9491 | 15.2738 | -0.01359 | -45.5398
170812 |  | 122-26rr |  | -40.9786 | 15.2803 | 0.0055659 | 102.0606 |  |  |  |  | -40.9479 | 15.2794 | -0.0084223 | -45.5282
170815 |  | 141-26 |  | -40.8006 | 71.287 | 0.056023 | -77.0673 |  |  |  |  | -40.6929 | 71.2868 | 0.0032782 | 8.3667
170815 |  | 141-26rr |  | -40.7996 | 71.285 | 0.052911 | -166.4472 |  |  |  |  | -40.6919 | 71.2849 | 0.00016822 | 8.3627
170815 |  | 141-26newref |  | -40.7909 | 71.3059 | 0.030816 | 337.3255 |  |  |  |  | -40.6832 | 71.3057 | -0.021934 | 8.3692
170816 |  | 112-26 |  | -40.7625 | -15.0084 | 0.011992 | 10.0245 |  |  |  |  | -40.774 | -15.0097 | 0.019367 | -74.3084
170822 |  | 131-26 |  | -40.7116 | 43.0507 | 0.0027172 | -112.4311 |  |  |  |  | -40.6429 | 43.0502 | -0.03045 | -18.669
170822 |  | 131-26rc | recleaned | -40.7014 | 43.0255 | 0.019833 | 227.1739 |  |  |  |  | -40.6326 | 43.025 | -0.013311 | -18.6663
170823 |  | 241-26rc | recleaned | -10.4082 | 71.3241 | 0.018639 | 242.5725 |  |  |  |  | -10.3364 | 71.3247 | -0.016227 | 39.1686
170823 |  | 121-26 | see recleaned? | -40.6527 | 16.1111 | 0.063666 | -99.2764 |  |  |  |  | -40.6212 | 16.1102 | 0.049283 | -44.364
170824 |  | 121-26rc | recleaned | -40.6418 | 16.1862 | 0.060361 | 384.6678 |  |  |  |  | -40.6101 | 16.1853 | 0.045932 | -44.2847
170825 |  | 211-26 |  | -10.422 | -15.1857 | 0.010691 | -42.4284 |  |  |  |  | -10.4696 | -15.1864 | 0.034376 | -46.2114
170828 |  | 212-26 |  | -10.3712 | -14.9218 | 0.0025528 | 82.1805 |  |  |  |  | -10.4186 | -14.9225 | 0.026082 | -45.9117
170829 |  | 142-26rr |  | -40.6926 | 71.1114 | 0.020575 | 291.5513 |  |  |  |  | -40.5853 | 71.1113 | -0.031981 | 8.272
EG3b |  |  |  | 2010 parameter set processing values:
|  |  |  | d13C | d18O | D47 | m/z44 diff, avg |  | 13C-18O | d47 | linearity, 45, 46
170803 |  | 231-cleaned |  | -10.3146 | 43.0906 | 0.017642 | 151.4387 |  | -53.4052 | 11.3804 | 0.046
170807 |  | 242-26new ref |  | -10.4172 | 71.4193 | -0.00079293 | 56.9268 |  | -81.8365 | 39.2332 | 0.095
170811 | 12V | 232-26-r1 |  | -10.4198 | 43.368 | -0.022335 | 34.4056 |  | -53.7878 | 11.5101
170812 | 12V | 122-26 |  | -40.969 | 15.2973 | 0.015592 | 47.8218 |  | -56.2663 | -45.4931
170815 | 12V | 141-26 |  | -40.7882 | 71.1825 | 0.044872 | -92.171 |  | -111.9707 | 8.2677
170816 | 12V | 112-26 |  | -40.7566 | -14.9276 | 0.0055747 | -3.0941 |  | -25.829 | -74.2316
170822 |  | 131-26rc |  | -40.6887 | 42.9943 | -0.011022 | 281.5539 |  | -83.683 | -18.714
170823 |  | 121-26 |  | -40.6412 | 16.1326 | 0.034604 | -31.8909 |  | -56.7737 | -44.3601
170825 |  | 211-26 |  | -10.4185 | -15.0969 | -0.0048403 | -100.7641 |  | 4.6785 | -46.1354
EG3c |  |  |  | 2010 parameter set processing values:
|  |  |  | d13C | d18O | D47 | amplitude 44, diff
linearity, 45, 46
170807 | test | 242-26new ref |  | -10.4179 | 71.5468 | -0.010558 | 8.7331 |  |  |  | 0.022
170811 | test | 232-26-r1 |  | -10.4196 | 43.4165 | -0.033841 | -1.6971 |  |  |  | 0.049
170812 |  | 122-26 |  | -40.9749 | 15.2684 | 0.067343 | 95.7957
170815 |  | 141-26 |  | -40.7902 | 71.3027 | 0.029957 | -82.9317
170816 |  | 112-26 |  | -40.7833 | -15.0253 | 0.093168 | 53.3532
170822 |  | 131-26rc | recleaned | -40.709 | 43.0046 | 0.051501 | 278.6862
170823 |  | 121-26 | see recleaned? | -40.6386 | 16.1134 | 0.049094 | -26.966
170825 |  | 211-26 |  | -10.417 | -15.1771 | 0.032772 | -71.0743
170828 |  | 212-26 |  | -10.3807 | -14.9309 | 0.049288 | 107.6965
170826 |  | 221-26 | test run ok, bad PBL on some standard runs only | -10.5111 | 15.5184 | 0.034514 | 46.6662
170827 |  | 142-26 | test run ok, bad PBL on some standard runs only | -40.7473 | 71.2315 | 0.03312 | 40.3106

## datapoints-plots
password to unlock: | unlock | Generate data for given sets of δ13C VPDB and δ18O VSMOW and reference values (CO2 gas, Δ47 calculations)
More information in other tabs |  | User input in yellow (unlocked) |  |  |  | Light blue for calculations |  |  | Orange for Δ47 vs δ47 linear fit
Parameters: see "Parameters" tab |  | Dark blue-data imported from yellow user input cell |  |  |  | Light purple, results from "full model" tab
ID, VSMOW 18O/16O; VPDB 13C/12C, lambda, K | [6 empty] | slope: | -0.00011903
data generation | Δ47 determination** |  |  |  | from model |  | intercept: | -0.0042298 |  | from model
2010 IUPAC | 2010 IUPAC |  | δ13CVPDB | δ18OVSMOW | δ47 | Δ47 vs. ref | linear fit, matrix only | Δ47, slope corr. | measured: | calc δ13C | calc δ18O | calc δ17O, from δ18O | δ13C calc - δ13C init | (δ13C - δ18O)
0.0020052 | 0.0020052 | reference (from "Parameters" tab) | -3.61 | 24.99 | 0 | -1.1102e-13 |  |  | uncorrected, or just initial PBL readings | -3.61 | 24.99 | 13.1178 | 0 | -28.6
0.01118 | 0.01118 | single point w/modifications | -3.61 | 24.99 | -0.0097809 | -0.010041 | -0.0042287 | -0.010042 |  | -3.61 | 24.9902 | 13.118 | 3.4032e-05 | -28.6002
0.528 | 0.528 | w/out added modifications | -3.61 | 24.99 | -0.0097809 | -0.010041 | -0.0042287 | -0.010042 | in order: | -3.61 | 24.9902 | 13.118 | 3.4032e-05 | -28.6002
0.010225 | 0.010225 | data points matrix: | -10.3146 | 43.0906 | 11.3506 | -0.0051415 | -0.0055809 | -0.0037905 | 0.017642 | -10.3269 | 43.096 | 22.528 | -0.012298 | -53.423
calculated: 17O/16O; 18O/total; 13C/total |  | (input data points here) | -10.4172 | 71.4193 | 39.2165 | -0.0049831 | -0.0088979 | -0.00031502 | -0.00079293 | -10.4427 | 71.4332 | 37.1022 | -0.025514 | -81.8759
0.00038476 | 0.00038476 | m/z 44 differences below | -10.4198 | 43.368 | 11.5206 | -0.0050632 | -0.0056012 | -0.0036919 | -0.022335 | -10.4323 | 43.3735 | 22.6717 | -0.012488 | -53.8058
0.99762 | 0.99762 | D47 for fitting to right | -40.969 | 15.2973 | -45.509 | 0.018375 | 0.0011872 | 0.012958 | 0.015592 | -40.9862 | 15.2953 | 8.047 | -0.017136 | -56.2815
0.98894 | 0.98894 |  | -40.7882 | 71.1825 | 8.211 | 0.018706 | -0.0052072 | 0.019684 | 0.044872 | -40.8313 | 71.1967 | 36.9814 | -0.043103 | -112.0281
ref gas values: | offset from label |  | -40.7566 | -14.9276 | -74.2321 | 0.018138 | 0.0046062 | 0.0093019 | 0.0055747 | -40.7593 | -14.9376 | -7.9151 | -0.0026472 | -25.8216
-3.61 | 0 | -3.61 | -40.6887 | 42.9943 | -18.7094 | 0.018335 | -0.0020028 | 0.016108 | -0.011022 | -40.7187 | 43.0002 | 22.4784 | -0.029968 | -83.7189
24.99 | 0 | 24.99 | -40.6412 | 16.1326 | -44.3943 | 0.018117 | 0.0010545 | 0.012833 | 0.034604 | -40.6585 | 16.1308 | 8.4849 | -0.017336 | -56.7893
see 'baseline-D17-D47" tab for more info. |  |  | -10.4185 | -15.0969 | -46.1316 | -0.0048957 | 0.0012613 | -0.010387 | -0.0048403 | -10.4033 | -15.1074 | -8.0054 | 0.015166 | 4.7041
baselines: 45,46,47, 44 | single point: |  |  |  | -21.3983 | -0.012506 | -0.0016828 | -0.015054 |  | 0.013926 | -0.0066402 | -0.003506 | 0.013926 | 0.020566
0 | 0 |  |  |  | -21.3983 | -0.012506 | -0.0016828 | -0.015054 |  | 0.013926 | -0.0066402 | -0.003506 | 0.013926 | 0.020566
0 | 0 |  |  |  | -21.3983 | -0.012506 | -0.0016828 | -0.015054 |  | 0.013926 | -0.0066402 | -0.003506 | 0.013926 | 0.020566
0.33473 | 0 |  |  |  | -21.3983 | -0.012506 | -0.0016828 | -0.015054 |  | 0.013926 | -0.0066402 | -0.003506 | 0.013926 | 0.020566
0 | 0 |  |  |  | -21.3983 | -0.012506 | -0.0016828 | -0.015054 |  | 0.013926 | -0.0066402 | -0.003506 | 0.013926 | 0.020566
Δ47, ref and sample | single point, sample only | 0 |  |  | -21.3983 | -0.012506 | -0.0016828 | -0.015054 |  | 0.013926 | -0.0066402 | -0.003506 | 0.013926 | 0.020566
0 | 0 | 0 |  |  | -21.3983 | -0.012506 | -0.0016828 | -0.015054 |  | 0.013926 | -0.0066402 | -0.003506 | 0.013926 | 0.020566
-0.0097826 | 0 | 0 |  |  | -21.3983 | -0.012506 | -0.0016828 | -0.015054 |  | 0.013926 | -0.0066402 | -0.003506 | 0.013926 | 0.020566
Δ17, ref and sample | single point, sample only | (bottom) |  |  | -21.3983 | -0.012506 | -0.0016828 | -0.015054 |  | 0.013926 | -0.0066402 | -0.003506 | 0.013926 | 0.020566
0 | 0 |  |  |  | -21.3983 | -0.012506 | -0.0016828 | -0.015054 |  | 0.013926 | -0.0066402 | -0.003506 | 0.013926 | 0.020566
0 | 0 |  |  |  | -21.3983 | -0.012506 | -0.0016828 | -0.015054 |  | 0.013926 | -0.0066402 | -0.003506 | 0.013926 | 0.020566
Linearity: 45,46,47, 44 | for no change, use: | m/z 44 working-sample (avg):
-0.024987 | 0; value+ lin44 for input to keep "measured" value unchanged | 151.4387 | -10.3146 | 43.0906 | 11.3399 | 0.0053574 | -0.0055796 | 0.0067072 |  | -10.3337 | 43.0812 | 22.5204 | -0.019059 | -53.4149
0.024013 | 0; value+ lin44 for input to keep "measured" value unchanged | 56.9268 | -10.4172 | 71.4193 | 39.2124 | -0.0010018 | -0.0088974 | 0.0036658 |  | -10.4452 | 71.4274 | 37.0993 | -0.028052 | -81.8726
0 | 0; value+ lin44 for input to keep "measured" value unchanged | 34.4056 | -10.4198 | 43.368 | 11.5182 | -0.0026776 | -0.0056009 | -0.0013066 |  | -10.4338 | 43.3702 | 22.6699 | -0.014024 | -53.804
-0.070987 | 0; subract from fitted value to return measured value | 47.8218 | -40.969 | 15.2973 | -45.5122 | 0.021698 | 0.0011876 | 0.016281 |  | -40.9883 | 15.2908 | 8.0446 | -0.019243 | -56.2791
fractional isotopologue abs. abundances |  | -92.171 | -40.7882 | 71.1825 | 8.2175 | 0.012192 | -0.005208 | 0.01317 |  | -40.8273 | 71.2061 | 36.9862 | -0.039052 | -112.0333
FI 45, 46, 47 | for no change, use: | -3.0941 | -40.7566 | -14.9276 | -74.2319 | 0.017925 | 0.0046061 | 0.0090889 |  | -40.7591 | -14.9373 | -7.9149 | -0.0025107 | -25.8218
0.98667 | 1 | 281.5539 | -40.6887 | 42.9943 | -18.7287 | 0.038064 | -0.0020005 | 0.035835 |  | -40.7311 | 42.9726 | 22.4641 | -0.042358 | -83.7037
1 | 1 | -31.8909 | -40.6412 | 16.1326 | -44.3922 | 0.015901 | 0.0010543 | 0.010617 |  | -40.6571 | 16.1338 | 8.4865 | -0.015931 | -56.7909
1 | 1 | -100.7641 | -10.4185 | -15.0969 | -46.1249 | -0.011756 | 0.0012605 | -0.017247 |  | -10.3988 | -15.0982 | -8.0005 | 0.019677 | 4.6994
Sample and ref signal |  |  | 0 | 0 | -21.3983 | -0.012506 | -0.0016828 | -0.015054 |  | 0.013926 | -0.0066402 | -0.003506 | 0.013926 | 0.020566
offset: | for no change, use: |  | 0 | 0 | -21.3983 | -0.012506 | -0.0016828 | -0.015054 |  | 0.013926 | -0.0066402 | -0.003506 | 0.013926 | 0.020566
0 | 0 |  | 0 | 0 | -21.3983 | -0.012506 | -0.0016828 | -0.015054 |  | 0.013926 | -0.0066402 | -0.003506 | 0.013926 | 0.020566
target value: | recommended: |  | 0 | 0 | -21.3983 | -0.012506 | -0.0016828 | -0.015054 |  | 0.013926 | -0.0066402 | -0.003506 | 0.013926 | 0.020566
12000 | 16000 |  | 0 | 0 | -21.3983 | -0.012506 | -0.0016828 | -0.015054 |  | 0.013926 | -0.0066402 | -0.003506 | 0.013926 | 0.020566
relative amplication |  |  | 0 | 0 | -21.3983 | -0.012506 | -0.0016828 | -0.015054 |  | 0.013926 | -0.0066402 | -0.003506 | 0.013926 | 0.020566
m/z 45, 46, 47 |  |  | 0 | 0 | -21.3983 | -0.012506 | -0.0016828 | -0.015054 |  | 0.013926 | -0.0066402 | -0.003506 | 0.013926 | 0.020566
100 | 100 |  | 0 | 0 | -21.3983 | -0.012506 | -0.0016828 | -0.015054 |  | 0.013926 | -0.0066402 | -0.003506 | 0.013926 | 0.020566
333.3333 | 333.3333 |  | 0 | 0 | -21.3983 | -0.012506 | -0.0016828 | -0.015054 |  | 0.013926 | -0.0066402 | -0.003506 | 0.013926 | 0.020566
3333.3333 | 3333.3333 |  | 0 | 0 | -21.3983 | -0.012506 | -0.0016828 | -0.015054 |  | 0.013926 | -0.0066402 | -0.003506 | 0.013926 | 0.020566
|  |  | 0 | 0 | -21.3983 | -0.012506 | -0.0016828 | -0.015054 |  | 0.013926 | -0.0066402 | -0.003506 | 0.013926 | 0.020566
|  |  | evaluation: | note:adjust min, max and sum sample ranges below to incorporate only lines used by sample data
|  |  | difference from measured: |  |  | square of differences:
|  |  | top | bottom |  | top | bottom
|  |  | -0.022784 | -0.012285 |  | 0.0005191 | 0.00015092
|  |  | -0.0041901 | -0.00020885 |  | 1.7557e-05 | 4.3618e-08
|  |  | 0.017272 | 0.019658 |  | 0.00029833 | 0.00038643
|  |  | 0.0027832 | 0.0061061 |  | 7.7463e-06 | 3.7284e-05
|  |  | -0.026166 | -0.03268 |  | 0.00068466 | 0.001068
|  |  | 0.012563 | 0.01235 |  | 0.00015783 | 0.00015253
|  |  | 0.029357 | 0.049086 |  | 0.00086186 | 0.0024094
|  |  | -0.016487 | -0.018703 |  | 0.00027181 | 0.00034979
|  |  | -5.5415e-05 | -0.0069158 |  | 3.0709e-09 | 4.7829e-05
|  |  | -0.012506 | -0.012506 |  | 0.00015641 | 0.00015641
|  |  | -0.012506 | -0.012506 |  | 0.00015641 | 0.00015641
|  |  | -0.012506 | -0.012506 |  | 0.00015641 | 0.00015641
|  |  | -0.012506 | -0.012506 |  | 0.00015641 | 0.00015641
|  |  | -0.012506 | -0.012506 |  | 0.00015641 | 0.00015641
|  |  | -0.012506 | -0.012506 |  | 0.00015641 | 0.00015641
|  |  | -0.012506 | -0.012506 |  | 0.00015641 | 0.00015641
|  |  | -0.012506 | -0.012506 |  | 0.00015641 | 0.00015641
|  |  | -0.012506 | -0.012506 |  | 0.00015641 | 0.00015641
|  |  | -0.012506 | -0.012506 |  | 0.00015641 | 0.00015641
Fitting |  |  | -0.012506 | -0.012506 |  | 0.00015641 | 0.00015641
Minimize sum of squares |  | min: | -0.026166 | -0.03268 | sum: | 0.0045394 | 0.0063228 | all to bottom
"Bottom section", i.e. H76, |  | max: | 0.029357 | 0.049086 |  | 0.0021342
takes into account |  | diff: | 0.055523 | 0.081766 |  |  | Results from previous fittings for datapoints from EG3b
m/z 44 amplitude differences |  | avg: | -0.00085621 | 0.001823 |  |  | (sum of squares after minimization)
|  | stdev: | 0.018749 | 0.023907 |  |  | To fit, use "Solver" in Data (install if needed), select target cell and cells with fitting parameters, and allow for negative values
** |  |  |  | 0.013692 |  |  | 0.0046023 | all EG3b
Δ47 determination** | [6 empty] | 0.0035343 | 2010fit
Use the settings used to process the data | [6 empty] | 0.0011248 | both, 2010, 1995
change "data generation" to find best fit | [6 empty] | 0.0021928 | 1995fit
[7 empty] | 0.0031845 | EG3a outliers
When fitting, start with measured linearity values | [6 empty] | 0.00077503 | all three possible outliers
for m/z 45 and 46 | [7 empty] | EG3c
[7 empty] | 0.0049151 | no outliers
Treat working gas as "0" Δ47 | [6 empty] | 0.004915 | Pierce criteria outliers (just 1)
Fit to baseline 47, PBL47, and Δ47 | [6 empty] | 0.0024684 | EG3a outliers
Use PBL corrected results, but still expect | [6 empty] | 0.0024683 | EG3a plus local Pierce
to have residual PBL47. | [6 empty] | 0.0019552 | EG3b-good set only
Expect Δ47 to be close to working gas value, but | [6 empty] | 0.0019551 | EG3b good only plus local Pierce
do not know actual value

## full model
Generate data for given sets of d13C VPDB and d18O VSMOW and reference values (CO2 gas, D47 calculations) | [9 empty] | password to unlock:
Green for input from "Interactive" tab |  |  |  |  | light blue for initial calculations, e.g. parameters |  |  |  |  | unlock | [8 empty] | Linearity and relative isotopologue absolute abundances
Yellow for user input (standard input from "interactive" tab) |  |  |  |  | purple for D47 results, also shown in "Interactive" | [5 empty] | model signals: |  | with offset | interactive | [6 empty] | linearity: on respective m/z scale |  | absolute abundances
|  | setup | [6 empty] | User Interactive--see Interactive page |  | max signal (44): | 12000 | 12000 | 12000 | value if combine all signals together | [5 empty] | per mil per mV (divide input by 1000) |  | relative to major isotopologue
parameters: |  | initial | solve-D47 |  | SSH | IUPAC-2010 | IsoDAT3.0 | Gonfiantini | 2010 IUPAC | 2010 IUPAC | rel amp, 45 | 100 |  | 100 | relative amplification with 44 amplification = 1, e.g. 3e8/3e8; for 45 3e10=300e8 so 300/3 or 100, etc. |  |  |  |  | std intens. | (default is 0) | same for single point |  | (default is 1)
| vsmow18 | 0.0020052 | 0.0020052 |  | 0.0020052 | 0.0020052 | 0.0020052 | 0.0020052 | 0.0020052 | 0.0020052 | rel amp, 46 | 333.3333 |  | 333.3333 | offset (m/z) from working gas: |  |  |  | m/z 45: | 14064.6521 | -2.0981e-05 | -2.0981e-05 |  | 0.98667
| vpdb13 | 0.01118 | 0.01118 |  | 0.000402 | 0.000385 | 0.000402 | 0.00038 | 0.01118 | 0.01118 | rel amp, 47 | 3333.3333 |  | 3333.3333 |  | 0 |  |  | m/z 46: | 16216.8857 | 1.7488e-05 | 1.7488e-05 |  | 1
| lambda | 0.528 | 0.528 |  | 0.011237 | 0.011803 | 0.011803 | 0.011237 | 0.528 | 0.528 |  |  |  | all: | single point modifications: |  |  |  | m/z 47: | 1866.1525 | 0 | 0 |  | 1
| K | 0.010225 | 0.010225 |  | 0.516 | 0.528 | 0.516 | 0.5164 | 0.010225 | 0.010225 | backgnd, 45 | 0 | 0 | 0 | 0 | sample and reference |  |  | m/z 44: | 11809.9697 | -7.0987e-05 | -7.0987e-05 |  | not applicable
[11 empty] | backgnd, 46 | 0 | 0 | 0 | 0 | sample and reference | [25 empty] | 0.0024063 |  | 0.0024063
| vsmow17 | 0.00038476 | 0.00038476 |  | 0.009924 | 0.010226 | 0.009924 | 0.009394 | 0.00038476 | 0.00038476 | backgnd, 47 | 0.33473 | 0.33473 | 0.33473 | 0 | sample and reference | [17 empty] | uses init R18= R46/2; both starting points give same result within 3 cycles
| vsmow16 | 0.99762 | 0.99762 |  | 0.9976 | 0.99762 | 0.9976 | 0.99762 | 0.99762 | 0.99762 | backgnd, 44 | 0 | 0 | 0 | 0 | sample and reference | [17 empty] | can use init R18=((1 +init18O)/1000)*VSMOW, with init18O=starting (given) 18O value
| vpdb12 | 0.98894 | 0.98894 |  | 0.98889 | 0.98833 | 0.98833 | 0.98889 | 0.98894 | 0.98894 | model D47, ref | 0 |  | 0 |  | use 0 for perfect stochastic distribution, note calculations are normalized to reference, so final reference value will be 0--but other values will shift | [8 empty] | Normalize values to ref:
[11 empty] | model D47, samp | -0.0097826 | -0.0097826 | -0.0097826 | 0 | applied as part per thousand shift on calc signal, before background adjustment | [8 empty] | below: use "solve-reported" parameters, in red | [8 empty] | minimize evaluation parameter: -3*R17^2+2*R45*R17+2*R18-R46; should equal 0 | [21 empty] | scrambled 47
interactive: | ref: | 13C | -3.61 | 18O | 24.99 | calc 17O | 13.1178 |  | Initial setup | Solve-D47 | Cap D17, reference | 0 |  | 0 |  | reference only | [8 empty] | here: use "solve-D47" parameters | [8 empty] | New R18 estimate from previous or init R18 - previous evaluation parameter/2.0022
| ref offset: | 13C | 0 | 18O | 0 | calc 17O | 13.1178 |  |  |  | Cap D17,sample | 0 | 0 | 0 | 0 | sample and single point only | [64 empty] | Should be "0" since 13C, 18O, D47 calc. w/ same parameters
from Interactive tab | init/ set values: |  | calculated values, use initial parameters, column "C": |  |  |  |  | fractional amounts: |  |  |  |  | sum, O | calculated CO2 values, fractional amounts: D47 corr |  |  |  |  | calc signals |  |  |  | R44 ratios from calculated signals: |  |  | ratios to reference: |  |  | theoretical reference R's from ref |  |  | working R's, RR*(R determination) |  |  | begin recalc of 13C and 18O, "solve" parameters |  |  |  |  | cycle1 |  |  |  | cycle2 |  |  |  | cycle3 |  |  |  | re-calc R13C | d relative to standard (first run), adjust to std values for 13C and 18O |  |  |  | calculated values: | [5 empty] | ratio to main isotope | [5 empty] | calculated CO2 values, from ratios |  |  |  | calculated R44 ratios from fractional space: |  |  | D47 | "measured" d47 from ref gas |  |  |  | can back out from D47, since excess 47 is loss of 45 and 46
| 13C | 18O | R13 | R18 | R17 | C12 | C16 | C13 | test13C | C18 | test18O | C17-R17hasD17 | frac. Amts | C47 | C44 | C45 | C46 | C47-D47 applied | m44 | m45 | m46 | m47 | R45 | R46 | R47 | RR45 | RR46 | RR47 | R13 | R18 | R17 | R45 | R46 | R47 | starting 18O | init R18 from R46 | init R17 | init R16 | init evaluation: | R18 | R17 | R16 | cycle 1 eval | R18 | R17 | R16 | cycle 2 eval | R18 | R17 | R16 | cycle 3 eval | calc R13 | d13C sa/ref | d18O sa/ref | d17O sa/ref--no correction to starting value | "initial calc R47" | R13 | R18 | R17 | C12 | C16 | C12/C12 =R12; C16/C16=R16 | C13 | re-calc-13C | C18 | re-calc18O | C17 | calc 17O--not normalized to ref gas | C44 | R45 | R46 | R47 | R45 | R46 | R47 | Δ47 vs. ref | δ47 | 13C | 18O | 17O | D45 | D46
reference: | -3.61 | 24.99 | 0.01114 | 0.0020553 | 0.0003898 | 0.98898 | 0.99756 | 0.011017 | -3.61 | 0.0020503 | 24.99 | 0.00038885 | 1 | 4.6645e-05 | 0.98416 | 0.011721 | 0.0040542 | 4.6645e-05 | 11809.9697 | 14064.6521 | 16216.8857 | 1866.1525 | 1.1909 | 1.3732 | 0.15802 | 1 | 1 | 1 | 0.01114 | 0.0020553 | 0.0003898 | 0.01192 | 0.0041195 | 4.7396e-05 | 24.99 | 0.0020597 | 0.00039025 | 0.99756 | 8.8462e-06 | 0.0020553 | 0.0003898 | 0.99756 | 2.1303e-10 | 0.0020553 | 0.0003898 | 0.99756 | 5.0003e-15 | 0.0020553 | 0.0003898 | 0.99756 | 0 | 0.01114 | -3.61 | 24.99 | 0 | 4.7396e-05 | 0.01114 | 0.0020553 | 0.0003898 | 0.98898 | 0.99756 | 1 | 0.011017 | -3.61 | 0.0020503 | 24.99 | 0.00038885 | 13.1178 | 0.98416 | 0.01192 | 0.0041195 | 4.7396e-05 | 0.01192 | 0.0041195 | 4.7396e-05 | -1.1102e-13 | 0 | -3.61 | 24.99 | 13.1178 | 0 | 0
single peak, modifications | -3.61 | 24.99 | 0.01114 | 0.0020553 | 0.0003898 | 0.98898 | 0.99756 | 0.011017 | -3.61 | 0.0020503 | 24.99 | 0.00038885 | 1 | 4.6645e-05 | 0.98416 | 0.011721 | 0.0040542 | 4.6645e-05 | 11809.9697 | 14064.6527 | 16216.8893 | 1866.1342 | 1.1909 | 1.3732 | 0.15801 | 1 | 1 | 0.99999 |  |  |  | 0.01192 | 0.0041195 | 4.7396e-05 | 24.99 | 0.0020597 | 0.00039025 | 0.99756 | 8.8462e-06 | 0.0020553 | 0.0003898 | 0.99756 | 2.1303e-10 | 0.0020553 | 0.0003898 | 0.99756 | 4.9995e-15 | 0.0020553 | 0.0003898 | 0.99756 | 0 | 0.01114 | -3.61 | 24.9902 | 0.00011766 | 4.7396e-05 | 0.01114 | 0.0020553 | 0.0003898 | 0.98898 | 0.99756 | 1 | 0.011017 | -3.61 | 0.0020503 | 24.9902 | 0.00038885 | 13.118 | 0.98416 | 0.01192 | 0.0041195 | 4.7396e-05 | 0.01192 | 0.0041195 | 4.7396e-05 | -0.010041 | -0.0097809 | -3.61 | 24.9902 | 13.118 | 0 | -2.2204e-13
without mod | -3.61 | 24.99 | 0.01114 | 0.0020553 | 0.0003898 | 0.98898 | 0.99756 | 0.011017 | -3.61 | 0.0020503 | 24.99 | 0.00038885 | 1 | 4.6645e-05 | 0.98416 | 0.011721 | 0.0040542 | 4.6645e-05 | 11809.9697 | 14064.6527 | 16216.8893 | 1866.1342 | 1.1909 | 1.3732 | 0.15801 | 1 | 1 | 0.99999 |  |  |  | 0.01192 | 0.0041195 | 4.7396e-05 | 24.99 | 0.0020597 | 0.00039025 | 0.99756 | 8.8462e-06 | 0.0020553 | 0.0003898 | 0.99756 | 2.1303e-10 | 0.0020553 | 0.0003898 | 0.99756 | 4.9995e-15 | 0.0020553 | 0.0003898 | 0.99756 | 0 | 0.01114 | -3.61 | 24.9902 | 0.00011766 | 4.7396e-05 | 0.01114 | 0.0020553 | 0.0003898 | 0.98898 | 0.99756 | 1 | 0.011017 | -3.61 | 0.0020503 | 24.9902 | 0.00038885 | 13.118 | 0.98416 | 0.01192 | 0.0041195 | 4.7396e-05 | 0.01192 | 0.0041195 | 4.7396e-05 | -0.010041 | -0.0097809 | -3.61 | 24.9902 | 13.118 | 0 | -2.2204e-13
| -10.3146 | 43.0906 | 0.011065 | 0.0020916 | 0.00039342 | 0.98906 | 0.99752 | 0.010944 | -10.3146 | 0.0020864 | 43.0906 | 0.00039245 | 1 | 4.7175e-05 | 0.98416 | 0.011654 | 0.0041257 | 4.7175e-05 | 11809.9048 | 13984.5077 | 16502.7414 | 1887.3241 | 1.1841 | 1.3974 | 0.15981 | 0.99431 | 1.0176 | 1.0114 |  |  |  | 0.011852 | 0.0041921 | 4.7934e-05 | 43.0906 | 0.002096 | 0.00039386 | 0.99752 | 8.8705e-06 | 0.0020916 | 0.00039342 | 0.99752 | 3.6124e-10 | 0.0020916 | 0.00039342 | 0.99752 | 1.4497e-14 | 0.0020916 | 0.00039342 | 0.99752 | 0 | 0.011065 | -10.3269 | 43.096 | 9.2884 | 4.7934e-05 | 0.011065 | 0.0020916 | 0.00039342 | 0.98906 | 0.99752 | 1 | 0.010944 | -10.3269 | 0.0020864 | 43.096 | 0.00039245 | 22.528 | 0.98416 | 0.011852 | 0.0041921 | 4.7934e-05 | 0.011852 | 0.0041921 | 4.7934e-05 | -0.0051415 | 11.3506 | -10.3269 | 43.096 | 22.528 | 0 | -2.2204e-13
| -10.4172 | 71.4193 | 0.011064 | 0.0021484 | 0.00039903 | 0.98906 | 0.99746 | 0.010943 | -10.4172 | 0.002143 | 71.4193 | 0.00039802 | 1 | 4.8469e-05 | 0.98404 | 0.011662 | 0.0042371 | 4.8469e-05 | 11808.4491 | 13994.4883 | 16948.5258 | 1939.0868 | 1.1851 | 1.4353 | 0.16421 | 0.99514 | 1.0453 | 1.0392 |  |  |  | 0.011862 | 0.0043059 | 4.9255e-05 | 71.4193 | 0.0021529 | 0.00039947 | 0.99745 | 8.998e-06 | 0.0021484 | 0.00039903 | 0.99746 | 4.9215e-10 | 0.0021484 | 0.00039903 | 0.99746 | 2.6632e-14 | 0.0021484 | 0.00039903 | 0.99746 | 0 | 0.011064 | -10.4427 | 71.4332 | 23.6738 | 4.9255e-05 | 0.011064 | 0.0021484 | 0.00039903 | 0.98906 | 0.99746 | 1 | 0.010942 | -10.4427 | 0.002143 | 71.4332 | 0.00039802 | 37.1022 | 0.98404 | 0.011862 | 0.0043059 | 4.9255e-05 | 0.011862 | 0.0043059 | 4.9255e-05 | -0.0049831 | 39.2165 | -10.4427 | 71.4332 | 37.1022 | 2.2204e-13 | -5.5511e-13
| -10.4198 | 43.368 | 0.011064 | 0.0020922 | 0.00039348 | 0.98906 | 0.99752 | 0.010943 | -10.4198 | 0.002087 | 43.368 | 0.0003925 | 1 | 4.7183e-05 | 0.98416 | 0.011653 | 0.0041268 | 4.7183e-05 | 11809.9042 | 13983.2472 | 16507.1229 | 1887.6412 | 1.184 | 1.3977 | 0.15984 | 0.99422 | 1.0179 | 1.0115 |  |  |  | 0.011851 | 0.0041932 | 4.7942e-05 | 43.368 | 0.0020966 | 0.00039392 | 0.99752 | 8.8708e-06 | 0.0020922 | 0.00039348 | 0.99752 | 3.6352e-10 | 0.0020922 | 0.00039348 | 0.99752 | 1.4681e-14 | 0.0020922 | 0.00039348 | 0.99752 | 0 | 0.011064 | -10.4323 | 43.3735 | 9.4301 | 4.7942e-05 | 0.011064 | 0.0020922 | 0.00039348 | 0.98906 | 0.99752 | 1 | 0.010943 | -10.4323 | 0.002087 | 43.3735 | 0.0003925 | 22.6717 | 0.98416 | 0.011851 | 0.0041932 | 4.7942e-05 | 0.011851 | 0.0041932 | 4.7942e-05 | -0.0050632 | 11.5206 | -10.4323 | 43.3735 | 22.6717 | 0 | 0
| -40.969 | 15.2973 | 0.010722 | 0.0020359 | 0.00038785 | 0.98939 | 0.99758 | 0.010608 | -40.969 | 0.002031 | 15.2973 | 0.00038692 | 1 | 4.4543e-05 | 0.98461 | 0.011311 | 0.0040174 | 4.4543e-05 | 11815.3498 | 13573.1484 | 16069.6948 | 1782.0372 | 1.1488 | 1.3601 | 0.15082 | 0.96461 | 0.99047 | 0.95449 |  |  |  | 0.011498 | 0.0040802 | 4.5239e-05 | 15.2973 | 0.0020401 | 0.00038828 | 0.99758 | 8.4764e-06 | 0.0020359 | 0.00038785 | 0.99758 | 5.1706e-10 | 0.0020359 | 0.00038785 | 0.99758 | 3.1245e-14 | 0.0020359 | 0.00038785 | 0.99758 | 0 | 0.010722 | -40.9862 | 15.2953 | -5.0052 | 4.5238e-05 | 0.010722 | 0.0020359 | 0.00038785 | 0.98939 | 0.99758 | 1 | 0.010608 | -40.9862 | 0.0020309 | 15.2953 | 0.00038691 | 8.047 | 0.98461 | 0.011498 | 0.0040802 | 4.5238e-05 | 0.011498 | 0.0040802 | 4.5238e-05 | 0.018375 | -45.509 | -40.9862 | 15.2953 | 8.047 | 0 | 0
| -40.7882 | 71.1825 | 0.010724 | 0.0021479 | 0.00039898 | 0.98939 | 0.99746 | 0.01061 | -40.7882 | 0.0021425 | 71.1825 | 0.00039797 | 1 | 4.7039e-05 | 0.98437 | 0.011332 | 0.0042373 | 4.7038e-05 | 11812.4252 | 13598.1095 | 16949.4275 | 1881.8666 | 1.1512 | 1.4349 | 0.15931 | 0.96663 | 1.045 | 1.0082 |  |  |  | 0.011522 | 0.0043046 | 4.7785e-05 | 71.1825 | 0.0021523 | 0.00039941 | 0.99745 | 8.7253e-06 | 0.002148 | 0.00039899 | 0.99746 | 7.6637e-10 | 0.002148 | 0.00039899 | 0.99746 | 6.6895e-14 | 0.002148 | 0.00039899 | 0.99746 | 0 | 0.010724 | -40.8313 | 71.1967 | 23.5546 | 4.7784e-05 | 0.010724 | 0.002148 | 0.00039899 | 0.98939 | 0.99746 | 1 | 0.01061 | -40.8313 | 0.0021425 | 71.1967 | 0.00039797 | 36.9814 | 0.98437 | 0.011522 | 0.0043046 | 4.7784e-05 | 0.011522 | 0.0043046 | 4.7784e-05 | 0.018706 | 8.211 | -40.8313 | 71.1967 | 36.9814 | 2.2204e-13 | -1.2212e-12
| -40.7566 | -14.9276 | 0.010725 | 0.0019753 | 0.00038171 | 0.98939 | 0.99765 | 0.010611 | -40.7566 | 0.0019706 | -14.9276 | 0.00038082 | 1 | 4.3208e-05 | 0.98474 | 0.011303 | 0.0038985 | 4.3208e-05 | 11816.8944 | 13563.4157 | 15593.6809 | 1728.6371 | 1.1478 | 1.3196 | 0.14629 | 0.9638 | 0.96101 | 0.92577 |  |  |  | 0.011488 | 0.0039588 | 4.3878e-05 | -14.9276 | 0.0019794 | 0.00038214 | 0.99764 | 8.3419e-06 | 0.0019752 | 0.00038171 | 0.99765 | 3.7714e-10 | 0.0019752 | 0.00038171 | 0.99765 | 1.6829e-14 | 0.0019752 | 0.00038171 | 0.99765 | 0 | 0.010725 | -40.7593 | -14.9376 | -20.7606 | 4.3877e-05 | 0.010725 | 0.0019752 | 0.00038171 | 0.98939 | 0.99765 | 1 | 0.010611 | -40.7593 | 0.0019706 | -14.9376 | 0.00038081 | -7.9151 | 0.98474 | 0.011488 | 0.0039588 | 4.3877e-05 | 0.011488 | 0.0039588 | 4.3877e-05 | 0.018138 | -74.2321 | -40.7593 | -14.9376 | -7.9151 | 0 | 0
| -40.6887 | 42.9943 | 0.010725 | 0.0020914 | 0.0003934 | 0.98939 | 0.99752 | 0.010612 | -40.6887 | 0.0020862 | 42.9943 | 0.00039243 | 1 | 4.5788e-05 | 0.98449 | 0.011323 | 0.0041264 | 4.5788e-05 | 11813.8745 | 13588.0911 | 16505.7199 | 1831.8434 | 1.1502 | 1.3971 | 0.15506 | 0.9658 | 1.0175 | 0.98129 |  |  |  | 0.011512 | 0.0041914 | 4.6509e-05 | 42.9943 | 0.0020957 | 0.00039383 | 0.99752 | 8.6021e-06 | 0.0020914 | 0.00039341 | 0.99752 | 6.3959e-10 | 0.0020914 | 0.00039341 | 0.99752 | 4.7197e-14 | 0.0020914 | 0.00039341 | 0.99752 | 0 | 0.010725 | -40.7187 | 43.0002 | 9.2394 | 4.6508e-05 | 0.010725 | 0.0020914 | 0.00039341 | 0.98939 | 0.99752 | 1 | 0.010611 | -40.7187 | 0.0020862 | 43.0002 | 0.00039243 | 22.4784 | 0.98449 | 0.011512 | 0.0041914 | 4.6508e-05 | 0.011512 | 0.0041914 | 4.6508e-05 | 0.018335 | -18.7094 | -40.7187 | 43.0002 | 22.4784 | 2.2204e-13 | -1.2212e-12
| -40.6412 | 16.1326 | 0.010726 | 0.0020375 | 0.00038802 | 0.98939 | 0.99758 | 0.010612 | -40.6412 | 0.0020326 | 16.1326 | 0.00038708 | 1 | 4.4595e-05 | 0.98461 | 0.011315 | 0.0040207 | 4.4594e-05 | 11815.2636 | 13577.7715 | 16082.7994 | 1784.1053 | 1.1492 | 1.3612 | 0.151 | 0.96495 | 0.99129 | 0.95561 |  |  |  | 0.011502 | 0.0040836 | 4.5292e-05 | 16.1326 | 0.0020418 | 0.00038845 | 0.99758 | 8.483e-06 | 0.0020375 | 0.00038802 | 0.99758 | 5.1791e-10 | 0.0020375 | 0.00038802 | 0.99758 | 3.1322e-14 | 0.0020375 | 0.00038802 | 0.99758 | 0 | 0.010726 | -40.6585 | 16.1308 | -4.573 | 4.5291e-05 | 0.010726 | 0.0020375 | 0.00038802 | 0.98939 | 0.99758 | 1 | 0.010612 | -40.6585 | 0.0020326 | 16.1308 | 0.00038708 | 8.4849 | 0.98461 | 0.011502 | 0.0040836 | 4.5291e-05 | 0.011502 | 0.0040836 | 4.5291e-05 | 0.018117 | -44.3943 | -40.6585 | 16.1308 | 8.4849 | -1.1102e-13 | -8.8818e-13
| -10.4185 | -15.0969 | 0.011064 | 0.0019749 | 0.00038168 | 0.98906 | 0.99765 | 0.010943 | -10.4185 | 0.0019703 | -15.0969 | 0.00038078 | 1 | 4.4505e-05 | 0.98441 | 0.011633 | 0.0038967 | 4.4504e-05 | 11812.9422 | 13959.3651 | 15586.8016 | 1780.512 | 1.1817 | 1.3195 | 0.15073 | 0.99226 | 0.9609 | 0.95387 |  |  |  | 0.011827 | 0.0039584 | 4.521e-05 | -15.0969 | 0.0019792 | 0.00038211 | 0.99764 | 8.6007e-06 | 0.0019749 | 0.00038168 | 0.99765 | 9.0846e-11 | 0.0019749 | 0.00038168 | 0.99765 | 9.0292e-16 | 0.0019749 | 0.00038168 | 0.99765 | 0 | 0.011064 | -10.4033 | -15.1074 | -20.8497 | 4.521e-05 | 0.011064 | 0.0019749 | 0.00038168 | 0.98906 | 0.99765 | 1 | 0.010943 | -10.4033 | 0.0019703 | -15.1074 | 0.00038078 | -8.0054 | 0.98441 | 0.011827 | 0.0039584 | 4.521e-05 | 0.011827 | 0.0039584 | 4.521e-05 | -0.0048957 | -46.1316 | -10.4033 | -15.1074 | -8.0054 | 2.2204e-13 | 0
| 0 | 0 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 0.011057 | -1.1102e-13 | 0.0020004 | 0 | 0.00038384 | 1 | 4.5651e-05 | 0.98423 | 0.011751 | 0.0039558 | 4.565e-05 | 11810.7973 | 14101.5342 | 15823.0229 | 1826.348 | 1.194 | 1.3397 | 0.15463 | 1.0026 | 0.97564 | 0.9786 |  |  |  | 0.01195 | 0.0040191 | 4.6382e-05 | 0 | 0.0020096 | 0.0003852 | 0.99761 | 8.7611e-06 | 0.0020052 | 0.00038476 | 0.99762 | 6.0308e-11 | 0.0020052 | 0.00038476 | 0.99762 | 3.773e-16 | 0.0020052 | 0.00038476 | 0.99762 | 0 | 0.01118 | 0.013926 | -0.0066402 | -12.9515 | 4.6382e-05 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 1 | 0.011057 | 0.013926 | 0.0020004 | -0.0066402 | 0.00038384 | -0.003506 | 0.98423 | 0.01195 | 0.0040191 | 4.6382e-05 | 0.01195 | 0.0040191 | 4.6382e-05 | -0.012506 | -21.3983 | 0.013926 | -0.0066402 | -0.003506 | 2.2204e-13 | 2.2204e-13
| 0 | 0 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 0.011057 | -1.1102e-13 | 0.0020004 | 0 | 0.00038384 | 1 | 4.5651e-05 | 0.98423 | 0.011751 | 0.0039558 | 4.565e-05 | 11810.7973 | 14101.5342 | 15823.0229 | 1826.348 | 1.194 | 1.3397 | 0.15463 | 1.0026 | 0.97564 | 0.9786 |  |  |  | 0.01195 | 0.0040191 | 4.6382e-05 | 0 | 0.0020096 | 0.0003852 | 0.99761 | 8.7611e-06 | 0.0020052 | 0.00038476 | 0.99762 | 6.0308e-11 | 0.0020052 | 0.00038476 | 0.99762 | 3.773e-16 | 0.0020052 | 0.00038476 | 0.99762 | 0 | 0.01118 | 0.013926 | -0.0066402 | -12.9515 | 4.6382e-05 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 1 | 0.011057 | 0.013926 | 0.0020004 | -0.0066402 | 0.00038384 | -0.003506 | 0.98423 | 0.01195 | 0.0040191 | 4.6382e-05 | 0.01195 | 0.0040191 | 4.6382e-05 | -0.012506 | -21.3983 | 0.013926 | -0.0066402 | -0.003506 | 2.2204e-13 | 2.2204e-13
| 0 | 0 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 0.011057 | -1.1102e-13 | 0.0020004 | 0 | 0.00038384 | 1 | 4.5651e-05 | 0.98423 | 0.011751 | 0.0039558 | 4.565e-05 | 11810.7973 | 14101.5342 | 15823.0229 | 1826.348 | 1.194 | 1.3397 | 0.15463 | 1.0026 | 0.97564 | 0.9786 |  |  |  | 0.01195 | 0.0040191 | 4.6382e-05 | 0 | 0.0020096 | 0.0003852 | 0.99761 | 8.7611e-06 | 0.0020052 | 0.00038476 | 0.99762 | 6.0308e-11 | 0.0020052 | 0.00038476 | 0.99762 | 3.773e-16 | 0.0020052 | 0.00038476 | 0.99762 | 0 | 0.01118 | 0.013926 | -0.0066402 | -12.9515 | 4.6382e-05 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 1 | 0.011057 | 0.013926 | 0.0020004 | -0.0066402 | 0.00038384 | -0.003506 | 0.98423 | 0.01195 | 0.0040191 | 4.6382e-05 | 0.01195 | 0.0040191 | 4.6382e-05 | -0.012506 | -21.3983 | 0.013926 | -0.0066402 | -0.003506 | 2.2204e-13 | 2.2204e-13
| 0 | 0 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 0.011057 | -1.1102e-13 | 0.0020004 | 0 | 0.00038384 | 1 | 4.5651e-05 | 0.98423 | 0.011751 | 0.0039558 | 4.565e-05 | 11810.7973 | 14101.5342 | 15823.0229 | 1826.348 | 1.194 | 1.3397 | 0.15463 | 1.0026 | 0.97564 | 0.9786 |  |  |  | 0.01195 | 0.0040191 | 4.6382e-05 | 0 | 0.0020096 | 0.0003852 | 0.99761 | 8.7611e-06 | 0.0020052 | 0.00038476 | 0.99762 | 6.0308e-11 | 0.0020052 | 0.00038476 | 0.99762 | 3.773e-16 | 0.0020052 | 0.00038476 | 0.99762 | 0 | 0.01118 | 0.013926 | -0.0066402 | -12.9515 | 4.6382e-05 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 1 | 0.011057 | 0.013926 | 0.0020004 | -0.0066402 | 0.00038384 | -0.003506 | 0.98423 | 0.01195 | 0.0040191 | 4.6382e-05 | 0.01195 | 0.0040191 | 4.6382e-05 | -0.012506 | -21.3983 | 0.013926 | -0.0066402 | -0.003506 | 2.2204e-13 | 2.2204e-13
| 0 | 0 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 0.011057 | -1.1102e-13 | 0.0020004 | 0 | 0.00038384 | 1 | 4.5651e-05 | 0.98423 | 0.011751 | 0.0039558 | 4.565e-05 | 11810.7973 | 14101.5342 | 15823.0229 | 1826.348 | 1.194 | 1.3397 | 0.15463 | 1.0026 | 0.97564 | 0.9786 |  |  |  | 0.01195 | 0.0040191 | 4.6382e-05 | 0 | 0.0020096 | 0.0003852 | 0.99761 | 8.7611e-06 | 0.0020052 | 0.00038476 | 0.99762 | 6.0308e-11 | 0.0020052 | 0.00038476 | 0.99762 | 3.773e-16 | 0.0020052 | 0.00038476 | 0.99762 | 0 | 0.01118 | 0.013926 | -0.0066402 | -12.9515 | 4.6382e-05 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 1 | 0.011057 | 0.013926 | 0.0020004 | -0.0066402 | 0.00038384 | -0.003506 | 0.98423 | 0.01195 | 0.0040191 | 4.6382e-05 | 0.01195 | 0.0040191 | 4.6382e-05 | -0.012506 | -21.3983 | 0.013926 | -0.0066402 | -0.003506 | 2.2204e-13 | 2.2204e-13
| 0 | 0 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 0.011057 | -1.1102e-13 | 0.0020004 | 0 | 0.00038384 | 1 | 4.5651e-05 | 0.98423 | 0.011751 | 0.0039558 | 4.565e-05 | 11810.7973 | 14101.5342 | 15823.0229 | 1826.348 | 1.194 | 1.3397 | 0.15463 | 1.0026 | 0.97564 | 0.9786 |  |  |  | 0.01195 | 0.0040191 | 4.6382e-05 | 0 | 0.0020096 | 0.0003852 | 0.99761 | 8.7611e-06 | 0.0020052 | 0.00038476 | 0.99762 | 6.0308e-11 | 0.0020052 | 0.00038476 | 0.99762 | 3.773e-16 | 0.0020052 | 0.00038476 | 0.99762 | 0 | 0.01118 | 0.013926 | -0.0066402 | -12.9515 | 4.6382e-05 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 1 | 0.011057 | 0.013926 | 0.0020004 | -0.0066402 | 0.00038384 | -0.003506 | 0.98423 | 0.01195 | 0.0040191 | 4.6382e-05 | 0.01195 | 0.0040191 | 4.6382e-05 | -0.012506 | -21.3983 | 0.013926 | -0.0066402 | -0.003506 | 2.2204e-13 | 2.2204e-13
| 0 | 0 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 0.011057 | -1.1102e-13 | 0.0020004 | 0 | 0.00038384 | 1 | 4.5651e-05 | 0.98423 | 0.011751 | 0.0039558 | 4.565e-05 | 11810.7973 | 14101.5342 | 15823.0229 | 1826.348 | 1.194 | 1.3397 | 0.15463 | 1.0026 | 0.97564 | 0.9786 |  |  |  | 0.01195 | 0.0040191 | 4.6382e-05 | 0 | 0.0020096 | 0.0003852 | 0.99761 | 8.7611e-06 | 0.0020052 | 0.00038476 | 0.99762 | 6.0308e-11 | 0.0020052 | 0.00038476 | 0.99762 | 3.773e-16 | 0.0020052 | 0.00038476 | 0.99762 | 0 | 0.01118 | 0.013926 | -0.0066402 | -12.9515 | 4.6382e-05 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 1 | 0.011057 | 0.013926 | 0.0020004 | -0.0066402 | 0.00038384 | -0.003506 | 0.98423 | 0.01195 | 0.0040191 | 4.6382e-05 | 0.01195 | 0.0040191 | 4.6382e-05 | -0.012506 | -21.3983 | 0.013926 | -0.0066402 | -0.003506 | 2.2204e-13 | 2.2204e-13
| 0 | 0 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 0.011057 | -1.1102e-13 | 0.0020004 | 0 | 0.00038384 | 1 | 4.5651e-05 | 0.98423 | 0.011751 | 0.0039558 | 4.565e-05 | 11810.7973 | 14101.5342 | 15823.0229 | 1826.348 | 1.194 | 1.3397 | 0.15463 | 1.0026 | 0.97564 | 0.9786 |  |  |  | 0.01195 | 0.0040191 | 4.6382e-05 | 0 | 0.0020096 | 0.0003852 | 0.99761 | 8.7611e-06 | 0.0020052 | 0.00038476 | 0.99762 | 6.0308e-11 | 0.0020052 | 0.00038476 | 0.99762 | 3.773e-16 | 0.0020052 | 0.00038476 | 0.99762 | 0 | 0.01118 | 0.013926 | -0.0066402 | -12.9515 | 4.6382e-05 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 1 | 0.011057 | 0.013926 | 0.0020004 | -0.0066402 | 0.00038384 | -0.003506 | 0.98423 | 0.01195 | 0.0040191 | 4.6382e-05 | 0.01195 | 0.0040191 | 4.6382e-05 | -0.012506 | -21.3983 | 0.013926 | -0.0066402 | -0.003506 | 2.2204e-13 | 2.2204e-13
| 0 | 0 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 0.011057 | -1.1102e-13 | 0.0020004 | 0 | 0.00038384 | 1 | 4.5651e-05 | 0.98423 | 0.011751 | 0.0039558 | 4.565e-05 | 11810.7973 | 14101.5342 | 15823.0229 | 1826.348 | 1.194 | 1.3397 | 0.15463 | 1.0026 | 0.97564 | 0.9786 |  |  |  | 0.01195 | 0.0040191 | 4.6382e-05 | 0 | 0.0020096 | 0.0003852 | 0.99761 | 8.7611e-06 | 0.0020052 | 0.00038476 | 0.99762 | 6.0308e-11 | 0.0020052 | 0.00038476 | 0.99762 | 3.773e-16 | 0.0020052 | 0.00038476 | 0.99762 | 0 | 0.01118 | 0.013926 | -0.0066402 | -12.9515 | 4.6382e-05 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 2 | 0.011057 | 0.013926 | 0.0020004 | -0.0066402 | 0.00038384 | -0.003506 | 0.98423 | 0.0478 | 0.016059 | 9.2763e-05 | 0.01195 | 0.0040191 | 4.6382e-05 | -0.012506 | -21.3983 | 0.013926 | -0.0066402 | -0.003506 | 2.2204e-13 | 2.2204e-13
| 0 | 0 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 0.011057 | -1.1102e-13 | 0.0020004 | 0 | 0.00038384 | 1 | 4.5651e-05 | 0.98423 | 0.011751 | 0.0039558 | 4.565e-05 | 11810.7973 | 14101.5342 | 15823.0229 | 1826.348 | 1.194 | 1.3397 | 0.15463 | 1.0026 | 0.97564 | 0.9786 |  |  |  | 0.01195 | 0.0040191 | 4.6382e-05 | 0 | 0.0020096 | 0.0003852 | 0.99761 | 8.7611e-06 | 0.0020052 | 0.00038476 | 0.99762 | 6.0308e-11 | 0.0020052 | 0.00038476 | 0.99762 | 3.773e-16 | 0.0020052 | 0.00038476 | 0.99762 | 0 | 0.01118 | 0.013926 | -0.0066402 | -12.9515 | 4.6382e-05 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 3 | 0.011057 | 0.013926 | 0.0020004 | -0.0066402 | 0.00038384 | -0.003506 | 0.98423 | 0.10755 | 0.03612 | 0.00013914 | 0.01195 | 0.0040191 | 4.6382e-05 | -0.012506 | -21.3983 | 0.013926 | -0.0066402 | -0.003506 | 2.2204e-13 | 2.2204e-13
| 0 | 0 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 0.011057 | -1.1102e-13 | 0.0020004 | 0 | 0.00038384 | 1 | 4.5651e-05 | 0.98423 | 0.011751 | 0.0039558 | 4.565e-05 | 11810.7973 | 14101.5342 | 15823.0229 | 1826.348 | 1.194 | 1.3397 | 0.15463 | 1.0026 | 0.97564 | 0.9786 |  |  |  | 0.01195 | 0.0040191 | 4.6382e-05 | 0 | 0.0020096 | 0.0003852 | 0.99761 | 8.7611e-06 | 0.0020052 | 0.00038476 | 0.99762 | 6.0308e-11 | 0.0020052 | 0.00038476 | 0.99762 | 3.773e-16 | 0.0020052 | 0.00038476 | 0.99762 | 0 | 0.01118 | 0.013926 | -0.0066402 | -12.9515 | 4.6382e-05 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 4 | 0.011057 | 0.013926 | 0.0020004 | -0.0066402 | 0.00038384 | -0.003506 | 0.98423 | 0.1912 | 0.064201 | 0.00018552 | 0.01195 | 0.0040191 | 4.6382e-05 | -0.012506 | -21.3983 | 0.013926 | -0.0066402 | -0.003506 | 2.2204e-13 | 2.2204e-13
44 offsets
151.4387 | -10.3146 | 43.0906 | 0.011065 | 0.0020916 | 0.00039342 | 0.98906 | 0.99752 | 0.010944 | -10.3146 | 0.0020864 | 43.0906 | 0.00039245 | 1 | 4.7175e-05 | 0.98416 | 0.011654 | 0.0041257 | 4.7175e-05 | 11660.9884 | 13808.0758 | 16294.4192 | 1863.5062 | 1.1841 | 1.3973 | 0.15981 | 0.9943 | 1.0176 | 1.0113 |  |  |  | 0.011852 | 0.004192 | 4.7933e-05 | 43.0906 | 0.002096 | 0.00039386 | 0.99752 | 8.8704e-06 | 0.0020916 | 0.00039342 | 0.99752 | 3.6124e-10 | 0.0020916 | 0.00039342 | 0.99752 | 1.4496e-14 | 0.0020916 | 0.00039342 | 0.99752 | 0 | 0.011065 | -10.3337 | 43.0812 | 9.2808 | 4.7933e-05 | 0.011065 | 0.0020916 | 0.00039342 | 0.98906 | 0.99752 | 1 | 0.010944 | -10.3337 | 0.0020864 | 43.0812 | 0.00039245 | 22.5204 | 0.98416 | 0.011852 | 0.004192 | 4.7933e-05 | 0.011852 | 0.004192 | 4.7933e-05 | 0.0053574 | 11.3399 | -10.3337 | 43.0812 | 22.5204 | 0 | 2.2204e-13
56.9268 | -10.4172 | 71.4193 | 0.011064 | 0.0021484 | 0.00039903 | 0.98906 | 0.99746 | 0.010943 | -10.4172 | 0.002143 | 71.4193 | 0.00039802 | 1 | 4.8469e-05 | 0.98404 | 0.011662 | 0.0042371 | 4.8469e-05 | 11752.4777 | 13928.1192 | 16868.0999 | 1929.888 | 1.1851 | 1.4353 | 0.16421 | 0.99514 | 1.0452 | 1.0392 |  |  |  | 0.011862 | 0.0043058 | 4.9255e-05 | 71.4193 | 0.0021529 | 0.00039947 | 0.99745 | 8.998e-06 | 0.0021484 | 0.00039903 | 0.99746 | 4.9215e-10 | 0.0021484 | 0.00039903 | 0.99746 | 2.6632e-14 | 0.0021484 | 0.00039903 | 0.99746 | 0 | 0.011063 | -10.4452 | 71.4274 | 23.6709 | 4.9255e-05 | 0.011063 | 0.0021484 | 0.00039903 | 0.98906 | 0.99746 | 1 | 0.010942 | -10.4452 | 0.002143 | 71.4274 | 0.00039802 | 37.0993 | 0.98404 | 0.011862 | 0.0043058 | 4.9255e-05 | 0.011862 | 0.0043058 | 4.9255e-05 | -0.0010018 | 39.2124 | -10.4452 | 71.4274 | 37.0993 | 2.2204e-13 | -4.4409e-13
34.4056 | -10.4198 | 43.368 | 0.011064 | 0.0020922 | 0.00039348 | 0.98906 | 0.99752 | 0.010943 | -10.4198 | 0.002087 | 43.368 | 0.0003925 | 1 | 4.7183e-05 | 0.98416 | 0.011653 | 0.0041268 | 4.7183e-05 | 11776.0719 | 13943.1672 | 16459.7812 | 1882.2291 | 1.184 | 1.3977 | 0.15984 | 0.99422 | 1.0179 | 1.0115 |  |  |  | 0.011851 | 0.0041932 | 4.7942e-05 | 43.368 | 0.0020966 | 0.00039392 | 0.99752 | 8.8708e-06 | 0.0020922 | 0.00039348 | 0.99752 | 3.6352e-10 | 0.0020922 | 0.00039348 | 0.99752 | 1.4681e-14 | 0.0020922 | 0.00039348 | 0.99752 | 0 | 0.011064 | -10.4338 | 43.3702 | 9.4284 | 4.7942e-05 | 0.011064 | 0.0020922 | 0.00039348 | 0.98906 | 0.99752 | 1 | 0.010943 | -10.4338 | 0.002087 | 43.3702 | 0.0003925 | 22.6699 | 0.98416 | 0.011851 | 0.0041932 | 4.7942e-05 | 0.011851 | 0.0041932 | 4.7942e-05 | -0.0026776 | 11.5182 | -10.4338 | 43.3702 | 22.6699 | 2.2204e-13 | -2.2204e-13
47.8218 | -40.969 | 15.2973 | 0.010722 | 0.0020359 | 0.00038785 | 0.98939 | 0.99758 | 0.010608 | -40.969 | 0.002031 | 15.2973 | 0.00038692 | 1 | 4.4543e-05 | 0.98461 | 0.011311 | 0.0040174 | 4.4543e-05 | 11768.3033 | 13519.0727 | 16005.6368 | 1774.9355 | 1.1488 | 1.3601 | 0.15082 | 0.96461 | 0.99047 | 0.95449 |  |  |  | 0.011498 | 0.0040802 | 4.5239e-05 | 15.2973 | 0.0020401 | 0.00038828 | 0.99758 | 8.4763e-06 | 0.0020359 | 0.00038785 | 0.99758 | 5.1706e-10 | 0.0020359 | 0.00038785 | 0.99758 | 3.1244e-14 | 0.0020359 | 0.00038785 | 0.99758 | 0 | 0.010722 | -40.9883 | 15.2908 | -5.0075 | 4.5238e-05 | 0.010722 | 0.0020359 | 0.00038785 | 0.98939 | 0.99758 | 1 | 0.010608 | -40.9883 | 0.0020309 | 15.2908 | 0.00038691 | 8.0446 | 0.98461 | 0.011498 | 0.0040802 | 4.5238e-05 | 0.011498 | 0.0040802 | 4.5238e-05 | 0.021698 | -45.5122 | -40.9883 | 15.2908 | 8.0446 | 2.2204e-13 | -6.6613e-13
-92.171 | -40.7882 | 71.1825 | 0.010724 | 0.0021479 | 0.00039898 | 0.98939 | 0.99746 | 0.01061 | -40.7882 | 0.0021425 | 71.1825 | 0.00039797 | 1 | 4.7039e-05 | 0.98437 | 0.011332 | 0.0042373 | 4.7038e-05 | 11903.0788 | 13702.5255 | 17079.6536 | 1896.3211 | 1.1512 | 1.4349 | 0.15931 | 0.96663 | 1.045 | 1.0082 |  |  |  | 0.011522 | 0.0043047 | 4.7785e-05 | 71.1825 | 0.0021523 | 0.00039941 | 0.99745 | 8.7253e-06 | 0.002148 | 0.00039899 | 0.99746 | 7.6638e-10 | 0.002148 | 0.00039899 | 0.99746 | 6.6895e-14 | 0.002148 | 0.00039899 | 0.99746 | 0 | 0.010724 | -40.8273 | 71.2061 | 23.5593 | 4.7785e-05 | 0.010724 | 0.002148 | 0.00039899 | 0.98939 | 0.99746 | 1 | 0.01061 | -40.8273 | 0.0021425 | 71.2061 | 0.00039797 | 36.9862 | 0.98437 | 0.011522 | 0.0043047 | 4.7785e-05 | 0.011522 | 0.0043047 | 4.7785e-05 | 0.012192 | 8.2175 | -40.8273 | 71.2061 | 36.9862 | 2.2204e-13 | -1.6653e-12
-3.0941 | -40.7566 | -14.9276 | 0.010725 | 0.0019753 | 0.00038171 | 0.98939 | 0.99765 | 0.010611 | -40.7566 | 0.0019706 | -14.9276 | 0.00038082 | 1 | 4.3208e-05 | 0.98474 | 0.011303 | 0.0038985 | 4.3208e-05 | 11819.9387 | 13566.912 | 15597.7027 | 1729.0828 | 1.1478 | 1.3196 | 0.14629 | 0.9638 | 0.96101 | 0.92577 |  |  |  | 0.011488 | 0.0039588 | 4.3878e-05 | -14.9276 | 0.0019794 | 0.00038214 | 0.99764 | 8.3419e-06 | 0.0019752 | 0.00038171 | 0.99765 | 3.7714e-10 | 0.0019752 | 0.00038171 | 0.99765 | 1.6829e-14 | 0.0019752 | 0.00038171 | 0.99765 | 0 | 0.010725 | -40.7591 | -14.9373 | -20.7604 | 4.3877e-05 | 0.010725 | 0.0019752 | 0.00038171 | 0.98939 | 0.99765 | 1 | 0.010611 | -40.7591 | 0.0019706 | -14.9373 | 0.00038081 | -7.9149 | 0.98474 | 0.011488 | 0.0039588 | 4.3877e-05 | 0.011488 | 0.0039588 | 4.3877e-05 | 0.017925 | -74.2319 | -40.7591 | -14.9373 | -7.9149 | 2.2204e-13 | -2.2204e-13
281.5539 | -40.6887 | 42.9943 | 0.010725 | 0.0020914 | 0.0003934 | 0.98939 | 0.99752 | 0.010612 | -40.6887 | 0.0020862 | 42.9943 | 0.00039243 | 1 | 4.5788e-05 | 0.98449 | 0.011323 | 0.0041264 | 4.5788e-05 | 11536.9146 | 13269.3649 | 16118.3399 | 1788.8632 | 1.1502 | 1.3971 | 0.15506 | 0.96578 | 1.0174 | 0.98127 |  |  |  | 0.011512 | 0.0041913 | 4.6508e-05 | 42.9943 | 0.0020957 | 0.00039383 | 0.99752 | 8.6019e-06 | 0.0020914 | 0.0003934 | 0.99752 | 6.3958e-10 | 0.0020914 | 0.0003934 | 0.99752 | 4.7197e-14 | 0.0020914 | 0.0003934 | 0.99752 | 0 | 0.010725 | -40.7311 | 42.9726 | 9.2253 | 4.6507e-05 | 0.010725 | 0.0020914 | 0.0003934 | 0.98939 | 0.99752 | 1 | 0.010611 | -40.7311 | 0.0020862 | 42.9726 | 0.00039242 | 22.4641 | 0.98449 | 0.011512 | 0.0041913 | 4.6507e-05 | 0.011512 | 0.0041913 | 4.6507e-05 | 0.038064 | -18.7287 | -40.7311 | 42.9726 | 22.4641 | 0 | -7.7716e-13
-31.8909 | -40.6412 | 16.1326 | 0.010726 | 0.0020375 | 0.00038802 | 0.98939 | 0.99758 | 0.010612 | -40.6412 | 0.0020326 | 16.1326 | 0.00038708 | 1 | 4.4595e-05 | 0.98461 | 0.011315 | 0.0040207 | 4.4594e-05 | 11846.6372 | 13613.8452 | 16125.5528 | 1788.8467 | 1.1492 | 1.3612 | 0.151 | 0.96495 | 0.99129 | 0.95561 |  |  |  | 0.011502 | 0.0040836 | 4.5292e-05 | 16.1326 | 0.0020418 | 0.00038845 | 0.99758 | 8.483e-06 | 0.0020376 | 0.00038802 | 0.99758 | 5.1791e-10 | 0.0020376 | 0.00038802 | 0.99758 | 3.1323e-14 | 0.0020376 | 0.00038802 | 0.99758 | 0 | 0.010726 | -40.6571 | 16.1338 | -4.5714 | 4.5291e-05 | 0.010726 | 0.0020376 | 0.00038802 | 0.98939 | 0.99758 | 1 | 0.010612 | -40.6571 | 0.0020326 | 16.1338 | 0.00038708 | 8.4865 | 0.98461 | 0.011502 | 0.0040836 | 4.5291e-05 | 0.011502 | 0.0040836 | 4.5291e-05 | 0.015901 | -44.3922 | -40.6571 | 16.1338 | 8.4865 | 2.2204e-13 | -4.4409e-13
-100.7641 | -10.4185 | -15.0969 | 0.011064 | 0.0019749 | 0.00038168 | 0.98906 | 0.99765 | 0.010943 | -10.4185 | 0.0019703 | -15.0969 | 0.00038078 | 1 | 4.4505e-05 | 0.98441 | 0.011633 | 0.0038967 | 4.4504e-05 | 11912.0517 | 14076.5473 | 15717.72 | 1795.4629 | 1.1817 | 1.3195 | 0.15073 | 0.99227 | 0.96091 | 0.95388 |  |  |  | 0.011827 | 0.0039584 | 4.521e-05 | -15.0969 | 0.0019792 | 0.00038212 | 0.99764 | 8.6008e-06 | 0.0019749 | 0.00038168 | 0.99765 | 9.0846e-11 | 0.0019749 | 0.00038168 | 0.99765 | 9.0206e-16 | 0.0019749 | 0.00038168 | 0.99765 | 0 | 0.011064 | -10.3988 | -15.0982 | -20.8449 | 4.521e-05 | 0.011064 | 0.0019749 | 0.00038168 | 0.98906 | 0.99765 | 1 | 0.010943 | -10.3988 | 0.0019703 | -15.0982 | 0.00038078 | -8.0005 | 0.98441 | 0.011827 | 0.0039584 | 4.521e-05 | 0.011827 | 0.0039584 | 4.521e-05 | -0.011756 | -46.1249 | -10.3988 | -15.0982 | -8.0005 | 0 | 0
0 | 0 | 0 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 0.011057 | -1.1102e-13 | 0.0020004 | 0 | 0.00038384 | 1 | 4.5651e-05 | 0.98423 | 0.011751 | 0.0039558 | 4.565e-05 | 11810.7973 | 14101.5342 | 15823.0229 | 1826.348 | 1.194 | 1.3397 | 0.15463 | 1.0026 | 0.97564 | 0.9786 |  |  |  | 0.01195 | 0.0040191 | 4.6382e-05 | 0 | 0.0020096 | 0.0003852 | 0.99761 | 8.7611e-06 | 0.0020052 | 0.00038476 | 0.99762 | 6.0308e-11 | 0.0020052 | 0.00038476 | 0.99762 | 3.773e-16 | 0.0020052 | 0.00038476 | 0.99762 | 0 | 0.01118 | 0.013926 | -0.0066402 | -12.9515 | 4.6382e-05 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 1 | 0.011057 | 0.013926 | 0.0020004 | -0.0066402 | 0.00038384 | -0.003506 | 0.98423 | 0.01195 | 0.0040191 | 4.6382e-05 | 0.01195 | 0.0040191 | 4.6382e-05 | -0.012506 | -21.3983 | 0.013926 | -0.0066402 | -0.003506 | 2.2204e-13 | 2.2204e-13
0 | 0 | 0 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 0.011057 | -1.1102e-13 | 0.0020004 | 0 | 0.00038384 | 1 | 4.5651e-05 | 0.98423 | 0.011751 | 0.0039558 | 4.565e-05 | 11810.7973 | 14101.5342 | 15823.0229 | 1826.348 | 1.194 | 1.3397 | 0.15463 | 1.0026 | 0.97564 | 0.9786 |  |  |  | 0.01195 | 0.0040191 | 4.6382e-05 | 0 | 0.0020096 | 0.0003852 | 0.99761 | 8.7611e-06 | 0.0020052 | 0.00038476 | 0.99762 | 6.0308e-11 | 0.0020052 | 0.00038476 | 0.99762 | 3.773e-16 | 0.0020052 | 0.00038476 | 0.99762 | 0 | 0.01118 | 0.013926 | -0.0066402 | -12.9515 | 4.6382e-05 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 1 | 0.011057 | 0.013926 | 0.0020004 | -0.0066402 | 0.00038384 | -0.003506 | 0.98423 | 0.01195 | 0.0040191 | 4.6382e-05 | 0.01195 | 0.0040191 | 4.6382e-05 | -0.012506 | -21.3983 | 0.013926 | -0.0066402 | -0.003506 | 2.2204e-13 | 2.2204e-13
0 | 0 | 0 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 0.011057 | -1.1102e-13 | 0.0020004 | 0 | 0.00038384 | 1 | 4.5651e-05 | 0.98423 | 0.011751 | 0.0039558 | 4.565e-05 | 11810.7973 | 14101.5342 | 15823.0229 | 1826.348 | 1.194 | 1.3397 | 0.15463 | 1.0026 | 0.97564 | 0.9786 |  |  |  | 0.01195 | 0.0040191 | 4.6382e-05 | 0 | 0.0020096 | 0.0003852 | 0.99761 | 8.7611e-06 | 0.0020052 | 0.00038476 | 0.99762 | 6.0308e-11 | 0.0020052 | 0.00038476 | 0.99762 | 3.773e-16 | 0.0020052 | 0.00038476 | 0.99762 | 0 | 0.01118 | 0.013926 | -0.0066402 | -12.9515 | 4.6382e-05 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 1 | 0.011057 | 0.013926 | 0.0020004 | -0.0066402 | 0.00038384 | -0.003506 | 0.98423 | 0.01195 | 0.0040191 | 4.6382e-05 | 0.01195 | 0.0040191 | 4.6382e-05 | -0.012506 | -21.3983 | 0.013926 | -0.0066402 | -0.003506 | 2.2204e-13 | 2.2204e-13
0 | 0 | 0 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 0.011057 | -1.1102e-13 | 0.0020004 | 0 | 0.00038384 | 1 | 4.5651e-05 | 0.98423 | 0.011751 | 0.0039558 | 4.565e-05 | 11810.7973 | 14101.5342 | 15823.0229 | 1826.348 | 1.194 | 1.3397 | 0.15463 | 1.0026 | 0.97564 | 0.9786 |  |  |  | 0.01195 | 0.0040191 | 4.6382e-05 | 0 | 0.0020096 | 0.0003852 | 0.99761 | 8.7611e-06 | 0.0020052 | 0.00038476 | 0.99762 | 6.0308e-11 | 0.0020052 | 0.00038476 | 0.99762 | 3.773e-16 | 0.0020052 | 0.00038476 | 0.99762 | 0 | 0.01118 | 0.013926 | -0.0066402 | -12.9515 | 4.6382e-05 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 1 | 0.011057 | 0.013926 | 0.0020004 | -0.0066402 | 0.00038384 | -0.003506 | 0.98423 | 0.01195 | 0.0040191 | 4.6382e-05 | 0.01195 | 0.0040191 | 4.6382e-05 | -0.012506 | -21.3983 | 0.013926 | -0.0066402 | -0.003506 | 2.2204e-13 | 2.2204e-13
0 | 0 | 0 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 0.011057 | -1.1102e-13 | 0.0020004 | 0 | 0.00038384 | 1 | 4.5651e-05 | 0.98423 | 0.011751 | 0.0039558 | 4.565e-05 | 11810.7973 | 14101.5342 | 15823.0229 | 1826.348 | 1.194 | 1.3397 | 0.15463 | 1.0026 | 0.97564 | 0.9786 |  |  |  | 0.01195 | 0.0040191 | 4.6382e-05 | 0 | 0.0020096 | 0.0003852 | 0.99761 | 8.7611e-06 | 0.0020052 | 0.00038476 | 0.99762 | 6.0308e-11 | 0.0020052 | 0.00038476 | 0.99762 | 3.773e-16 | 0.0020052 | 0.00038476 | 0.99762 | 0 | 0.01118 | 0.013926 | -0.0066402 | -12.9515 | 4.6382e-05 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 1 | 0.011057 | 0.013926 | 0.0020004 | -0.0066402 | 0.00038384 | -0.003506 | 0.98423 | 0.01195 | 0.0040191 | 4.6382e-05 | 0.01195 | 0.0040191 | 4.6382e-05 | -0.012506 | -21.3983 | 0.013926 | -0.0066402 | -0.003506 | 2.2204e-13 | 2.2204e-13
0 | 0 | 0 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 0.011057 | -1.1102e-13 | 0.0020004 | 0 | 0.00038384 | 1 | 4.5651e-05 | 0.98423 | 0.011751 | 0.0039558 | 4.565e-05 | 11810.7973 | 14101.5342 | 15823.0229 | 1826.348 | 1.194 | 1.3397 | 0.15463 | 1.0026 | 0.97564 | 0.9786 |  |  |  | 0.01195 | 0.0040191 | 4.6382e-05 | 0 | 0.0020096 | 0.0003852 | 0.99761 | 8.7611e-06 | 0.0020052 | 0.00038476 | 0.99762 | 6.0308e-11 | 0.0020052 | 0.00038476 | 0.99762 | 3.773e-16 | 0.0020052 | 0.00038476 | 0.99762 | 0 | 0.01118 | 0.013926 | -0.0066402 | -12.9515 | 4.6382e-05 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 1 | 0.011057 | 0.013926 | 0.0020004 | -0.0066402 | 0.00038384 | -0.003506 | 0.98423 | 0.01195 | 0.0040191 | 4.6382e-05 | 0.01195 | 0.0040191 | 4.6382e-05 | -0.012506 | -21.3983 | 0.013926 | -0.0066402 | -0.003506 | 2.2204e-13 | 2.2204e-13
0 | 0 | 0 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 0.011057 | -1.1102e-13 | 0.0020004 | 0 | 0.00038384 | 1 | 4.5651e-05 | 0.98423 | 0.011751 | 0.0039558 | 4.565e-05 | 11810.7973 | 14101.5342 | 15823.0229 | 1826.348 | 1.194 | 1.3397 | 0.15463 | 1.0026 | 0.97564 | 0.9786 |  |  |  | 0.01195 | 0.0040191 | 4.6382e-05 | 0 | 0.0020096 | 0.0003852 | 0.99761 | 8.7611e-06 | 0.0020052 | 0.00038476 | 0.99762 | 6.0308e-11 | 0.0020052 | 0.00038476 | 0.99762 | 3.773e-16 | 0.0020052 | 0.00038476 | 0.99762 | 0 | 0.01118 | 0.013926 | -0.0066402 | -12.9515 | 4.6382e-05 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 1 | 0.011057 | 0.013926 | 0.0020004 | -0.0066402 | 0.00038384 | -0.003506 | 0.98423 | 0.01195 | 0.0040191 | 4.6382e-05 | 0.01195 | 0.0040191 | 4.6382e-05 | -0.012506 | -21.3983 | 0.013926 | -0.0066402 | -0.003506 | 2.2204e-13 | 2.2204e-13
0 | 0 | 0 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 0.011057 | -1.1102e-13 | 0.0020004 | 0 | 0.00038384 | 1 | 4.5651e-05 | 0.98423 | 0.011751 | 0.0039558 | 4.565e-05 | 11810.7973 | 14101.5342 | 15823.0229 | 1826.348 | 1.194 | 1.3397 | 0.15463 | 1.0026 | 0.97564 | 0.9786 |  |  |  | 0.01195 | 0.0040191 | 4.6382e-05 | 0 | 0.0020096 | 0.0003852 | 0.99761 | 8.7611e-06 | 0.0020052 | 0.00038476 | 0.99762 | 6.0308e-11 | 0.0020052 | 0.00038476 | 0.99762 | 3.773e-16 | 0.0020052 | 0.00038476 | 0.99762 | 0 | 0.01118 | 0.013926 | -0.0066402 | -12.9515 | 4.6382e-05 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 1 | 0.011057 | 0.013926 | 0.0020004 | -0.0066402 | 0.00038384 | -0.003506 | 0.98423 | 0.01195 | 0.0040191 | 4.6382e-05 | 0.01195 | 0.0040191 | 4.6382e-05 | -0.012506 | -21.3983 | 0.013926 | -0.0066402 | -0.003506 | 2.2204e-13 | 2.2204e-13
0 | 0 | 0 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 0.011057 | -1.1102e-13 | 0.0020004 | 0 | 0.00038384 | 1 | 4.5651e-05 | 0.98423 | 0.011751 | 0.0039558 | 4.565e-05 | 11810.7973 | 14101.5342 | 15823.0229 | 1826.348 | 1.194 | 1.3397 | 0.15463 | 1.0026 | 0.97564 | 0.9786 |  |  |  | 0.01195 | 0.0040191 | 4.6382e-05 | 0 | 0.0020096 | 0.0003852 | 0.99761 | 8.7611e-06 | 0.0020052 | 0.00038476 | 0.99762 | 6.0308e-11 | 0.0020052 | 0.00038476 | 0.99762 | 3.773e-16 | 0.0020052 | 0.00038476 | 0.99762 | 0 | 0.01118 | 0.013926 | -0.0066402 | -12.9515 | 4.6382e-05 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 2 | 0.011057 | 0.013926 | 0.0020004 | -0.0066402 | 0.00038384 | -0.003506 | 0.98423 | 0.0478 | 0.016059 | 9.2763e-05 | 0.01195 | 0.0040191 | 4.6382e-05 | -0.012506 | -21.3983 | 0.013926 | -0.0066402 | -0.003506 | 2.2204e-13 | 2.2204e-13
0 | 0 | 0 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 0.011057 | -1.1102e-13 | 0.0020004 | 0 | 0.00038384 | 1 | 4.5651e-05 | 0.98423 | 0.011751 | 0.0039558 | 4.565e-05 | 11810.7973 | 14101.5342 | 15823.0229 | 1826.348 | 1.194 | 1.3397 | 0.15463 | 1.0026 | 0.97564 | 0.9786 |  |  |  | 0.01195 | 0.0040191 | 4.6382e-05 | 0 | 0.0020096 | 0.0003852 | 0.99761 | 8.7611e-06 | 0.0020052 | 0.00038476 | 0.99762 | 6.0308e-11 | 0.0020052 | 0.00038476 | 0.99762 | 3.773e-16 | 0.0020052 | 0.00038476 | 0.99762 | 0 | 0.01118 | 0.013926 | -0.0066402 | -12.9515 | 4.6382e-05 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 3 | 0.011057 | 0.013926 | 0.0020004 | -0.0066402 | 0.00038384 | -0.003506 | 0.98423 | 0.10755 | 0.03612 | 0.00013914 | 0.01195 | 0.0040191 | 4.6382e-05 | -0.012506 | -21.3983 | 0.013926 | -0.0066402 | -0.003506 | 2.2204e-13 | 2.2204e-13
0 | 0 | 0 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 0.011057 | -1.1102e-13 | 0.0020004 | 0 | 0.00038384 | 1 | 4.5651e-05 | 0.98423 | 0.011751 | 0.0039558 | 4.565e-05 | 11810.7973 | 14101.5342 | 15823.0229 | 1826.348 | 1.194 | 1.3397 | 0.15463 | 1.0026 | 0.97564 | 0.9786 |  |  |  | 0.01195 | 0.0040191 | 4.6382e-05 | 0 | 0.0020096 | 0.0003852 | 0.99761 | 8.7611e-06 | 0.0020052 | 0.00038476 | 0.99762 | 6.0308e-11 | 0.0020052 | 0.00038476 | 0.99762 | 3.773e-16 | 0.0020052 | 0.00038476 | 0.99762 | 0 | 0.01118 | 0.013926 | -0.0066402 | -12.9515 | 4.6382e-05 | 0.01118 | 0.0020052 | 0.00038476 | 0.98894 | 0.99762 | 4 | 0.011057 | 0.013926 | 0.0020004 | -0.0066402 | 0.00038384 | -0.003506 | 0.98423 | 0.1912 | 0.064201 | 0.00018552 | 0.01195 | 0.0040191 | 4.6382e-05 | -0.012506 | -21.3983 | 0.013926 | -0.0066402 | -0.003506 | 2.2204e-13 | 2.2204e-13